## exp1
25mA current average |  |  | 0 deg
Turbidity | Voltage 1 | Voltage 2 | Voltage max
0 | 1.385 | 1.145 | 1.385
10.7 | 1.196 | 1.365 | 1.365
21.1 | 1.38 | 1.345 | 1.38
30.4 | 1.076 | 1.28 | 1.28
35.4 | 1.05 | 0.9 | 1.05
79.9 | 1.08 | 1.15 | 1.15
208 | 0.681 | 0.645 | 0.681
464 | 0.145 | 0.127 | 0.145
50mA constant current |  |  | 0 deg
Turbidity | Voltage 1 | Voltage 2 | Voltage max
0 | 2.4 | 2.8 | 2.8
10.7 | 2.539 | 2.9 | 2.9
21.1 | 2.9 | 2.9 | 2.9
30.4 | 2.322 | 2.8 | 2.8
35.4 | 2 | 3.2 | 3.2
79.9 | 2.45 | 2.27 | 2.45
208 | 1.7 | 1.6 | 1.7
464 | 0.317 | 0.301 | 0.317
25mA constant current |  |  | 0 deg
Turbidity | Voltage 1 | Voltage 2 | Voltage max
0 | 1.344 | 1.22 | 1.344
10.7 | 1.42 | 1.59 | 1.59
21.1 | 1.609 | 1.58 | 1.609
30.4 | 1.38 | 1.57 | 1.57
35.4 | 0.97 | 1.154 | 1.154
79.9 | 1.3 | 1.2 | 1.3
208 | 0.8 | 0.768 | 0.8
464 | 0.2 | 0.183 | 0.2
50mA constant current |  |  | 90 deg
Turbidity | Voltage 1 | Voltage 2 | Voltage max
0 | 4.5 | 3.992
10.7
21.1
30.4
35.4
79.9 | 4.209
208
464
1479
2791

## exp2
68.5mA
Turbidity | Transmission | 90 deg
24.1 | 3.19 | 0.108
54.6 | 2.823 | 0.136
86.3 | 2.756 | 0.163
133 | 2.139 | 0.192
216 | 1.92 | 0.262
288 | 1.781 | 0.344
390 | 1.727 | 0.405
1007 | 0.736 | 0.74
1288 | 0.205 | 1.325
5420 | 0.205 | 1.825

## exp3
~300mA | Outdoors
NTU | Transmission | Transmission DC | 90 deg | 90 deg DC | True transmission | True 90 deg
Air | 3.055 | 0.013 | 3.256 | 1.603 | 3.042 | 1.653
0 | 3.618 | 0.013 | 2.392 | 0.784 | 3.605 | 1.608
24.1 | 3.402 | 0.013 | 2.445 | 0.797 | 3.389 | 1.648
54.6 | 3.9 | 0.013 | 2.43 | 0.811 | 3.887 | 1.619
86.3 | 3.615 | 0.013 | 2.525 | 0.824 | 3.602 | 1.701
133 | 3.402 | 0.013 | 2.565 | 0.877 | 3.389 | 1.688
216 | 3.336 | 0.013 | 2.684 | 0.877 | 3.323 | 1.807
288 | 2.472 | 0.013 | 2.804 | 0.85 | 2.459 | 1.954
390 | 0.198 | 0.013 | 2.87 | 0.85 | 0.185 | 2.02
1007 | 1.103 | 0.013 | 3.2 | 0.811 | 1.09 | 2.389
~300mA | Outdoors
NTU | 90 deg | 90 deg DC | True 90 deg
Air | 2.247 | 0 | 2.247
0 | 1.626 | 0 | 1.626
24.1 | 1.661 | 0 | 1.661
54.6 | 1.681 | 0 | 1.681
86.3 | 1.73 | 0 | 1.73
133 | 1.759 | 0 | 1.759
216 | 1.863 | 0 | 1.863
288 | 1.926 | 0 | 1.926
390 | 2.08 | 0 | 2.08
1007 | 2.479 | 0 | 2.479

## Turbo_2.0
| 3.1k resistor feedback resistor | two LEDS | 100 samples average | Indoor | One resistor on the left and one on the right
NTU | On (V) | Off (V) | On-Off (V)
0 | 2389.21 | 495.98 | 1893.23
24.1 | 2462.82 | 495.84 | 1966.98
54.6 | 2503.85 | 496.3 | 2007.55
86.3 | 2551.34 | 496.38 | 2054.96
133 | 2664.45 | 496.09 | 2168.36
216 | 2729.21 | 496.08 | 2233.13
288 | 2821.16 | 496.61 | 2324.55
390 | 2943.11 | 495.89 | 2447.22
596 | 3292.93 | 496.02 | 2796.91
750 | 3857.64 | 496.82 | 3360.82
1007 | 3447.33 | 496.49 | 2950.84
1288 | 3839.7 | 496.68 | 3343.02
1838 | 5171.75 | 497.82 | 4673.93 |  |  | s
2539 | 4779.22 | 497.23 | 4281.99
3003 | 4813.09 | 497.11 | 4315.98
3517 | 4982.79 | 497.24 | 4485.55

## Turbo_2.1
NTU | On (V) | Off (V) | On-Off (V) | Inside | Two leds stacked on top
0 | 2052.34 | 496.1 | 1556.24 |  | 175.5681 | 1.7765 | -90838.8164 | -93.3954
24.1 | 2085.8 | 495.93 | 1589.87 |  | 175.5681 | 1.7765 | -90838.8164 | -63.7988
54.6 | 2108 | 496.18 | 1611.82 |  | 175.5681 | 1.7765 | -90838.8164 | 19.5325
86.3 | 2116.32 | 495.82 | 1620.5 |  | 175.5681 | 1.7765 | -90838.8164 | -28.8911
133 | 2157.22 | 496.14 | 1661.08 |  | 175.5681 | 1.7765 | -90838.8164 | 99.9513
216 | 2194.4 | 495.82 | 1698.58 |  | 175.5681 | 1.7765 | -90838.8164 | 109.8215
288 | 2270.99 | 496.15 | 1774.84 |  | 175.5681 | 1.7765 | -90838.8164 | 303.8245
390 | 2378.56 | 495.79 | 1882.77 |  | 175.5681 | 1.7765 | -90838.8164 | 431.7229
596 | 2512.35 | 496.3 | 2016.05 |  | 175.5681 | 1.7765 | -90838.8164 | 758.9466
839 | 2637.74 | 496.34 | 2141.4 |  | 175.5681 | 1.7765 | -90838.8164 | 988.7302
1007 | 2719.85 | 496.36 | 2223.49 |  | 175.5681 | 1.7765 | -90838.8164 | 1138.1137
1288 | 2830.49 | 496.28 | 2334.21 |  | 175.5681 | 1.7765 | -90838.8164 | 1320.6251
1827 | 3245.56 | 496.65 | 2748.91 |  | 175.5681 | 1.7765 | -90838.8164 | 2122.9757
2539 | 3462.93 | 497.25 | 2965.68 |  | 175.5681 | 1.7765 | -90838.8164 | 2614.4841
3003 | 3618.2 | 497.14 | 3121.06 |  | 175.5681 | 1.7765 | -90838.8164 | 2871.0157
3394 | 3753.09 | 497.03 | 3256.06 |  | 175.5681 | 1.7765 | -90838.8164 | 3091.3414
| Outside |  | 2023-03-16
NTU | On (V) | Off (V) | On-Off (V)
0 | 2285.43 | 653.24 | 1632.19 |  | -0.23325 | 1.9785 | -4550.502 | -181.2489
24.1 | 2323.29 | 643.44 | 1679.85 |  | -0.23325 | 1.9785 | -4550.502 | -104.0588
54.6 | 2340.47 | 640.66 | 1699.81 |  | -0.23325 | 1.9785 | -4550.502 | -69.4205
86.3 | 2422.05 | 643.03 | 1779.02 |  | -0.23325 | 1.9785 | -4550.502 | 91.429
133 | 2377.95 | 608.13 | 1769.82 |  | -0.23325 | 1.9785 | -4550.502 | 12.3195
216 | 2511.07 | 649.04 | 1862.03 |  | -0.23325 | 1.9785 | -4550.502 | 266.1492
288 | 2508.24 | 617.77 | 1890.47 |  | -0.23325 | 1.9785 | -4550.502 | 267.8438
390 | 2636.37 | 644.52 | 1991.85 |  | -0.23325 | 1.9785 | -4550.502 | 515.1038
596 | 2767.58 | 647.48 | 2120.1 |  | -0.23325 | 1.9785 | -4550.502 | 774.0064
839 | 2892.7 | 625.08 | 2267.62 |  | -0.23325 | 1.9785 | -4550.502 | 1026.7753
1007 | 2954.58 | 627.17 | 2327.41 |  | -0.23325 | 1.9785 | -4550.502 | 1148.7146
1288 | 3056.71 | 635.5 | 2421.21 |  | -0.23325 | 1.9785 | -4550.502 | 1348.8312
1827 | 3381.69 | 634.32 | 2747.37 |  | -0.23325 | 1.9785 | -4550.502 | 1992.0645
2539 | 3651.11 | 622.48 | 3028.63 |  | -0.23325 | 1.9785 | -4550.502 | 2527.8613
3003 | 3900.52 | 637.01 | 3263.51 |  | -0.23325 | 1.9785 | -4550.502 | 3017.9185
3394 | 3916.78 | 634.34 | 3282.44 |  | -0.23325 | 1.9785 | -4550.502 | 3050.711
Disabled measurement LED
388.66 | 181.21 | 207.45 | 2.32
519.92 | 520.42 | -0.5 | 1.8
181.98 | 182.19 | -0.21 | 1.96
182.83 | 183 | -0.17 | 1.67
495.06 | 495.63 | -0.57 | 1.88
291.85 | 292.14 | -0.29 | 2.18
167.2 | 167.09 | 0.11 | 1.99
187.4 | 187.41 | -0.01 | 1.91
135.11 | 135.38 | -0.27 | 1.76
185.81 | 185.7 | 0.11 | 2.82
126.05 | 126.22 | -0.17 | 1.78
414.91 | 414.98 | -0.07 | 2.37
191.67 | 194.05 | -2.38 | 6.15
572.13 | 572.6 | -0.47 | 1.94
573.92 | 574.83 | -0.91 | 2.47
493.49 | 493.95 | -0.46 | 2.09
Enable measurement LED but turn them away
126.05 | 126.17 | -0.12 | 2.02
126.05 | 126.4 | -0.35 | 1.61
131.15 | 126.44 | 4.71 | 1.97
130.73 | 126.53 | 4.2 | 1.63
365.84 | 361.22 | 4.62 | 1.98
376.38 | 371.69 | 4.69 | 1.86
397.75 | 393.26 | 4.49 | 1.68
343.34 | 338.37 | 4.97 | 1.88
310.71 | 305.78 | 4.93 | 2.04
380.7 | 376 | 4.7 | 2.21
445.45 | 440.81 | 4.64 | 2.13
526.85 | 521.74 | 5.11 | 2.14
652.64 | 647.95 | 4.69 | 1.89
668.29 | 663.46 | 4.83 | 1.84
715.43 | 710.71 | 4.72 | 2.1
310.49 | 306.26 | 4.23 | 1.9

## Turbo_3
Glued in the LEDS
High delay | Low delay | Sample len | Samples | Bits | Gain
2 | 20000 | 2 | 100 | 14 | Low gain
Inside with glued in parts
NTU | On (V) | Off (V) | On-Off (V) |  | 2.0185 | -3385.8422
0 | 2082.13 | 496.02 | 1586.11
24.1 | 2100.25 | 495.77 | 1604.48
54.6 | 2115.9 | 496.04 | 1619.86
86.3 | 2188.85 | 495.95 | 1692.9
133 | 2234.66 | 495.8 | 1738.86
216 | 2264.46 | 495.69 | 1768.77
288 | 2323.89 | 495.72 | 1828.17
390 | 2440.06 | 496.35 | 1943.71
596 | 2546.38 | 496.35 | 2050.03
839 | 2713.18 | 496.25 | 2216.93
1007 | 2743.01 | 496.33 | 2246.68
1288 | 2863.78 | 496.06 | 2367.72
1827 | 3244.4 | 496.07 | 2748.33
2539 | 3445.49 | 496.07 | 2949.42
3003 | 3581.62 | 496.26 | 3085.36
3394 | 3657.21 | 496.05 | 3161.16
High delay | Low delay | Sample len | Samples | Bits | Gain
0 | 20000 | 0 | 100 | 12 | Low gain
NTU | On (V) | Off (V) | On-Off (V) | SD
0 | 833.06 | 238.26 | 594.8 | 3.07 |  | 4.7743 | -2849.3961
24.1 | 805.26 | 215.39 | 589.87 | 3.03
54.6 | 869.53 | 264.65 | 604.88 | 2.7
86.3 | 878.7 | 257.77 | 620.93 | 4.54
133 | 922.49 | 281.87 | 640.62 | 3.35
216 | 894.35 | 244.18 | 650.17 | 4.02
288 | 854.77 | 208.27 | 646.5 | 4.14
390 | 911.59 | 224.46 | 687.13 | 3.94
596 | 909.2 | 192.13 | 717.07 | 3.19
839 | 942.88 | 176.44 | 766.44 | 4.53
Outside
High delay | Low delay | Sample len | Samples | Bits | Gain
0 | 20000 | 0 | 100 | 14 | Low gain
NTU | On (V) | Off (V) | On-Off (V) | SD
0 | 7099.41 | 5614.33 | 1485.08 | 111.16
24.1 | 7083.31 | 5593.11 | 1490.2 | 136.68
54.6 | 7084.95 | 5605.72 | 1479.23 | 125.66
86.3 | 7086.34 | 5614.83 | 1471.51 | 118.74
133 | 7115.41 | 5630.34 | 1485.07 | 128.04
216 | 7094.89 | 5618.67 | 1476.22 | 118.06
288 | 7075.31 | 5608.98 | 1466.33 | 123.42
390 | 7088.48 | 5628.95 | 1459.53 | 123.82
Repeatability test
High delay | Low delay | Sample len | Samples | Bits | Gain
0 | 20000 | 0 | 100 | 14 | Low gain
NTU | On (V) | Off (V) | On-Off (V) | SD |  | Factory calibration
0 | 1674.26 | 1285.84 | 388.42 | 58.29 | Inside | TRUE
0 | 1733.24 | 1327.83 | 405.41 | 69.47 | Inside | TRUE
0 | 1708.1 | 1314.03 | 394.07 | 69.16 | Inside | TRUE
0 | 1705.54 | 1309.77 | 395.77 | 40.97 | Inside | TRUE
0 | 1686.29 | 1300.05 | 386.24 | 62.09 | Inside | TRUE
0 | 1706.57 | 1302.99 | 403.58 | 47.53 | Outside | TRUE
0 | 1717.16 | 1308.4 | 408.76 | 45.38 | Outside | TRUE | Outside | 16:53:35.166 -> 3855.96,1256.71,2599.25,61.31
0 | 1703.54 | 1294.25 | 409.29 | 47.25 | Outside | TRUE | Inside | 16:53:53.561 -> 2729.89,495.82,2234.07,62.14
0 | 1673.67 | 1273.71 | 399.96 | 47.34 | Outside | TRUE
0 | 1657.35 | 1259.7 | 397.65 | 49.35 | Outside | TRUE
0 | 1642 | 1248.38 | 393.62 | 45 | Smothered with jacket | TRUE
0 | 624.73 | 124.34 | 500.39 | 3.03 | Storage room | TRUE
0 | 632.7 | 123.75 | 508.95 | 4.34 | Storage room | TRUE
0 | 637.67 | 124.38 | 513.29 | 5.12 | Storage room | TRUE
0 | 640.6 | 124.35 | 516.25 | 3.34 | Storage room | TRUE
0 | 643.63 | 124.1 | 519.53 | 3.71 | Storage room | TRUE
1007 | 2757.12 | 496.43 | 2260.69 | 62.37 | Inside | TRUE
1007 | 2761.44 | 496.71 | 2264.73 | 62.56 | Inside | TRUE
1007 | 2763.16 | 496.48 | 2266.68 | 62.61 | Inside | TRUE
1007 | 2761.37 | 496.15 | 2265.22 | 62.87 | Inside | TRUE
0 | 2094.54 | 496.69 | 1597.85 | 43.77 | Inside | TRUE
0 | 2092.72 | 496.64 | 1596.08 | 44 | Inside | TRUE
0 | 2090.75 | 496.29 | 1594.46 | 43.75 | Inside | TRUE
0 | 2091.39 | 496.38 | 1595.01 | 43.85 | Inside | TRUE
0 | 7069.18 | 5639.23 | 1429.95 | 104.18 | Outside | TRUE
0 | 2369.84 | 681.91 | 1687.93 | 43.97 | Outside | TRUE
0 | 2336.05 | 665.08 | 1670.97 | 43.11 | Outside | TRUE
0 | 2374.97 | 727.4 | 1647.57 | 310.13 | Outside | TRUE
0 | 2950.47 | 1169.57 | 1780.9 | 39.44 | Outside | TRUE
0 | 3223.35 | 1274.91 | 1948.44 | 42.78 | Outside | TRUE
0 | 7106.43 | 5637.11 | 1469.32 | 122.96 | Outside | TRUE
0 | 7107.65 | 5636.29 | 1471.36 | 103.91 | Outside | TRUE
0 | 7118.32 | 5643.78 | 1474.54 | 111.59 | Outside | TRUE
0 | 7132.96 | 5670.19 | 1462.77 | 104.63 | Outside | TRUE
0 | 7126.12 | 5661.74 | 1464.38 | 104.64 | Outside | TRUE
0 | 7090.82 | 5625.44 | 1465.38 | 104.19 | Outside | TRUE
1007 | 3559.92 | 1083.8 | 2476.12 | 62.3 | Outside | TRUE
1007 | 3748.03 | 1239.77 | 2508.26 | 58.93 | Outside | TRUE
1007 | 3691.47 | 1230.83 | 2460.64 | 54.71 | Outside | TRUE
1007 | 3665.12 | 1230.71 | 2434.41 | 58.41 | Outside | TRUE
1007 | 7127.63 | 5663.58 | 1464.05 | 122.75 | Outside | TRUE
1007 | 7107.1 | 5663.89 | 1443.21 | 99.19 | Outside | TRUE
1007 | 7135.53 | 5654.49 | 1481.04 | 119.56 | Outside | TRUE
1007 | 7120.17 | 5663.97 | 1456.2 | 114.54 | Outside | TRUE
1007 | 7152.01 | 5698.88 | 1453.13 | 125.26 | Outside | TRUE
1007 | 7123.63 | 5667.92 | 1455.71 | 112.34 | Outside | TRUE
1007 | 7104.79 | 5650.06 | 1454.73 | 133.27 | Outside | TRUE
1007 | 7126.94 | 5666.17 | 1460.77 | 109.4 | Outside | TRUE
0 | 7125.74 | 5649.01 | 1476.73 | 137.5 | Outside | TRUE
0 | 7160.94 | 5700.78 | 1460.16 | 111.73 | Outside | TRUE
0 | 7089 | 5646.07 | 1442.93 | 108.45 | Outside | TRUE
0 | 7128.03 | 5661.21 | 1466.82 | 117.6 | Outside | TRUE
0 | 7112.08 | 5638.42 | 1473.66 | 116.48 | Outside | TRUE
0 | 7109.53 | 5642.82 | 1466.71 | 101.05 | Outside | TRUE
0 | 7128.95 | 5666.29 | 1462.66 | 113.43 | Outside | TRUE
0 | 7115.84 | 5649.67 | 1466.17 | 111.29 | Outside | TRUE
Foil | 7110.8 | 5655.4 | 1455.4 | 125.85 | Outside | TRUE
Foil | 7122.66 | 5662.58 | 1460.08 | 130.4 | Outside | TRUE
Foil | 7106.2 | 5647.09 | 1459.11 | 118.19 | Outside | TRUE
Foil | 7115.53 | 5655.68 | 1459.85 | 116.05 | Outside | TRUE
Foil | 7098.36 | 5627.41 | 1470.95 | 104.94 | Outside | TRUE
Foil | 7140.93 | 5674.38 | 1466.55 | 114.79 | Outside | TRUE
Foil | 7114.79 | 5654.78 | 1460.01 | 123.42 | Outside | TRUE
Foil | 7096 | 5639 | 1457 | 122.71 | Outside | TRUE
Foil | 14963.27 | 1297.67 | 13665.6 | 249.44 | Outside | TRUE
Foil | 13497.3 | 846.89 | 12650.41 | 280.5 | Outside | TRUE
Foil | 13250.67 | 845.88 | 12404.79 | 260.34 | Outside | TRUE
Foil | 13117.66 | 844.06 | 12273.6 | 253.61 | Outside | TRUE
Foil | 13145.23 | 817.35 | 12327.88 | 266.95 | Outside | TRUE
Foil | 13148.84 | 815.65 | 12333.19 | 267.99 | Outside | TRUE
Foil | 13241.68 | 876.82 | 12364.86 | 268.8 | Outside | TRUE
Foil | 13410.57 | 834.2 | 12576.37 | 277.4 | Outside | TRUE
Reconnected diode |  |  | 2023-03-25
0 | 2140.42 | 603.24 | 1537.18 | 40.85 | Outside | TRUE
0 | 2115.28 | 600.21 | 1515.07 | 38.97 | Outside | TRUE
0 | 2088.41 | 599.54 | 1488.87 | 36.06 | Outside | TRUE
0 | 2070.49 | 599.5 | 1470.99 | 34.99 | Outside | TRUE
0 | 2059.19 | 599.2 | 1459.99 | 34.89 | Outside | TRUE
0 | 2050.5 | 599.07 | 1451.43 | 33.79 | Outside | TRUE
0 | 2043.45 | 598.17 | 1445.28 | 34.42 | Outside | TRUE
0 | 2035.45 | 597.87 | 1437.58 | 33.33 | Outside | TRUE
1007 | 2766.51 | 597.82 | 2168.69 | 58.93 | Outside | TRUE
1007 | 2710.18 | 598.33 | 2111.85 | 54.86 | Outside | TRUE
1007 | 2691.72 | 599.29 | 2092.43 | 53.96 | Outside | TRUE
1007 | 2653.59 | 596.94 | 2056.65 | 50.74 | Outside | TRUE
1007 | 2692.12 | 598.89 | 2093.23 | 55.56 | Outside | TRUE
1007 | 2650.93 | 598.75 | 2052.18 | 52.01 | Outside | TRUE
1007 | 2634.11 | 599.75 | 2034.36 | 50.45 | Outside | TRUE
0 | 2065.1 | 495.26 | 1569.84 | 42.1 | Storage room | TRUE
0 | 2024.34 | 495.33 | 1529.01 | 38.6 | Storage room | TRUE
0 | 2002.45 | 495.53 | 1506.92 | 35.89 | Storage room | TRUE
0 | 1986.37 | 495.26 | 1491.11 | 35.68 | Storage room | TRUE
0 | 1974.74 | 495.42 | 1479.32 | 34.91 | Storage room | TRUE
0 | 1967.58 | 495.07 | 1472.51 | 35.47 | Storage room | TRUE
0 | 1961.24 | 495.08 | 1466.16 | 35.12 | Storage room | TRUE
0 | 1953.95 | 495.61 | 1458.34 | 34.57 | Storage room | TRUE
1007 | 2730.58 | 495.64 | 2234.94 | 59.96 | Storage room | TRUE
1007 | 2664.15 | 495.65 | 2168.5 | 54.83 | Storage room | TRUE
1007 | 2629.62 | 495.28 | 2134.34 | 53.16 | Storage room | TRUE
1007 | 2599.25 | 495.45 | 2103.8 | 51.88 | Storage room | TRUE
1007 | 2571.18 | 495.14 | 2076.04 | 49.77 | Storage room | TRUE
1007 | 2557.48 | 495.4 | 2062.08 | 49.85 | Storage room | TRUE
1007 | 2555.4 | 494.96 | 2060.44 | 50.52 | Storage room | TRUE
1007 | 2551.59 | 495.18 | 2056.41 | 51.06 | Storage room | TRUE
Calibration curve | HIGH GAIN
0 | 6476.21 | 496.56 | 5979.65 | 180.18 | Storage room | TRUE
24.1 | 6550.43 | 497.2 | 6053.23 | 186.83 | Storage room | TRUE
54.6 | 6846.18 | 496.79 | 6349.39 | 194.47 | Storage room | TRUE
86.3 | 6937.71 | 497.33 | 6440.38 | 215.44 | Storage room | TRUE
133 | 7145.02 | 497.12 | 6647.9 | 219.42 | Storage room | TRUE
216 | 7383.14 | 497.2 | 6885.94 | 214.81 | Storage room | TRUE
288 | 7703.21 | 497.73 | 7205.48 | 232.35 | Storage room | TRUE
390 | 7886.25 | 497.14 | 7389.11 | 224.8 | Storage room | TRUE
596 | 8404.74 | 497.54 | 7907.2 | 234.56 | Storage room | TRUE
839 | 9131.33 | 497 | 8634.33 | 249.68 | Storage room | TRUE
1007 | 9365.72 | 498.04 | 8867.68 | 278.73 | Storage room | TRUE
1288 | 9693.06 | 498.07 | 9194.99 | 274.81 | Storage room | TRUE
1827 | 10988.17 | 498.11 | 10490.06 | 302.24 | Storage room | TRUE
2539 | 12083 | 498.57 | 11584.43 | 337.58 | Storage room | TRUE
3003 | 12949.3 | 500.06 | 12449.24 | 352.08 | Storage room | TRUE
3394 | 12951.47 | 498.68 | 12452.79 | 367.24 | Storage room | TRUE
0 | 6899.81 | 893.1 | 6006.71 | 71.83 | Outside | TRUE
24.1 | 7174 | 873.51 | 6300.49 | 178.93 | Outside | TRUE
54.6 | 6998.47 | 808.07 | 6190.4 | 108.34 | Outside | TRUE
86.3 | 7162.95 | 820.77 | 6342.18 | 183.64 | Outside | TRUE
133 | 7309.9 | 829.66 | 6480.24 | 200.75 | Outside | TRUE
216 | 7664.21 | 827.55 | 6836.66 | 207.1 | Outside | TRUE
288 | 7826.24 | 849.25 | 6976.99 | 249.77 | Outside | TRUE
390 | 8318.76 | 906.51 | 7412.25 | 213.35 | Outside | TRUE
596 | 8790.69 | 974.94 | 7815.75 | 222.12 | Outside | TRUE
839 | 9443.13 | 968.46 | 8474.67 | 240.77 | Outside | TRUE
1007 | 9626.93 | 958.75 | 8668.18 | 248.53 | Outside | TRUE
1288 | 10193.25 | 952.47 | 9240.78 | 271.28 | Outside | TRUE
1827 | 11754.65 | 982.03 | 10772.62 | 299.17 | Outside | TRUE
2539 | 12439.38 | 1012.73 | 11426.65 | 137.92 | Outside | TRUE
3003 | 12867.03 | 1046.81 | 11820.22 | 157.78 | Outside | TRUE
3394 | 12748.53 | 901.54 | 11846.99 | 312.41 | Outside | TRUE
Calibration curve | LOW GAIN |  |  | SD
0 | 2122.06 | 495.76 | 1626.3 | 44.91 | Storage room
24.1 | 2179.08 | 495.45 | 1683.63 | 48.51 | Storage room
54.6 | 2172.07 | 495.63 | 1676.44 | 47.74 | Storage room
86.3 | 2213.07 | 495.11 | 1717.96 | 47.97 | Storage room
133 | 2276.84 | 495.68 | 1781.16 | 50.88 | Storage room
216 | 2341.82 | 495.36 | 1846.46 | 53.61 | Storage room
288 | 2398.5 | 495.66 | 1902.84 | 54.77 | Storage room
390 | 2459 | 495.44 | 1963.56 | 57.44 | Storage room
596 | 2573.32 | 495.47 | 2077.85 | 60.27 | Storage room
839 | 2738.12 | 495.99 | 2242.13 | 61.59 | Storage room
1007 | 2812.07 | 495.58 | 2316.49 | 63.81 | Storage room
1288 | 2914.36 | 496.1 | 2418.26 | 64.04 | Storage room
1827 | 3338.68 | 495.81 | 2842.87 | 79.12 | Storage room
2539 | 3541.94 | 495.9 | 3046.04 | 86.74 | Storage room
3003 | 3605.58 | 495.53 | 3110.05 | 87.65 | Storage room
3394 | 3819.7 | 495.22 | 3324.48 | 90.75 | Storage room
0 | 2237.1 | 594.08 | 1643.02 | 43.69 | Outside
24.1 | 2280.92 | 591.84 | 1689.08 | 45.08 | Outside
54.6 | 2266.86 | 588.68 | 1678.18 | 39.29 | Outside
86.3 | 2343.11 | 590.42 | 1752.69 | 48.24 | Outside
133 | 2386.47 | 589.62 | 1796.85 | 51.8 | Outside
216 | 2474.44 | 590.16 | 1884.28 | 54.51 | Outside
288 | 2540.19 | 589.03 | 1951.16 | 59.91 | Outside
390 | 2601.93 | 589.34 | 2012.59 | 59.23 | Outside
596 | 2691.08 | 590.05 | 2101.03 | 57.45 | Outside
839 | 2884.38 | 590.65 | 2293.73 | 60.89 | Outside
1007 | 2901.75 | 591.4 | 2310.35 | 66.6 | Outside
1288 | 3033.44 | 591.41 | 2442.03 | 68.16 | Outside
1827 | 3322.73 | 591.75 | 2730.98 | 63.21 | Outside
2539 | 3471.72 | 592.51 | 2879.21 | 85.06 | Outside
3003 | 3818.41 | 592.22 | 3226.19 | 87.34 | Outside
3394 | 3910.84 | 591.81 | 3319.03 | 103.41 | Outside
9V battery
0 | 2679.36 | 789.32 | 1890.04 | 44.81 | Outside
0 | 2521.07 | 691.78 | 1829.29 | 48.14 | Outside
1007 | 3228.73 | 670.62 | 2558.11 | 68 | Outside
1007 | 3160.38 | 668.87 | 2491.51 | 62.57 | Outside
0 | 2266.12 | 495.16 | 1770.96 | 46 | Inside
0 | 2226.13 | 495.11 | 1731.02 | 43.52 | Inside
1007 | 2945.14 | 495.42 | 2449.72 | 66.46 | Inside
1007 | 2888.18 | 495.42 | 2392.76 | 60.15 | Inside
Fixed timing issue
Low gain | 90 degree | Inside |  |  | Rectify On | Rectify Off | Rectify Diff
0 | 2602.92 | 494.6 | 2108.32 | 6.83 | #NUM!
24.1 | 2619.15 | 494.89 | 2124.26 | 5.98
54.6 | 2670.16 | 494.45 | 2175.71 | 9.84
86.3 | 2750.92 | 495.03 | 2255.89 | 8.23
133 | 2796.07 | 494.89 | 2301.18 | 6.62
216 | 2949.52 | 495.14 | 2454.38 | 9.31
288 | 2958.38 | 495.11 | 2463.27 | 67.56
390 | 3072.96 | 495.25 | 2577.71 | 10.1
596 | 3219.07 | 494.89 | 2724.18 | 9.69
839 | 3421.88 | 494.92 | 2926.96 | 9.56
1007 | 3511.93 | 495.04 | 3016.89 | 10.43
1288 | 3674.68 | 494.86 | 3179.82 | 12.41
1827 | 4118.67 | 495.61 | 3623.06 | 12.29
2539 | 4508.94 | 495.37 | 4013.57 | 32.96
3003 | 4585.21 | 495.47 | 4089.74 | 15.57
3394 | 4823.7 | 495.53 | 4328.17 | 13.55
High gain | 90 degree Inside
0 | 8095.14 | 496.93 | 7598.21 | 26.14
24.1 | 8243.26 | 496.87 | 7746.39 | 26.79
54.6 | 8201.76 | 496.98 | 7704.78 | 26.91
86.3 | 8456.76 | 496.76 | 7960 | 28.55
133 | 8791.34 | 496.78 | 8294.56 | 30.38
216 | 9213.05 | 496.99 | 8716.06 | 35.29
288 | 9670.48 | 497.33 | 9173.15 | 45.1
390 | 9924.66 | 497.31 | 9427.35 | 39.37
596 | 10586.43 | 497.82 | 10088.61 | 37.84
839 | 11489.63 | 497.4 | 10992.23 | 37.14
1007 | 11794.04 | 497.38 | 11296.66 | 41.28
1288 | 12374.95 | 498.04 | 11876.91 | 41.78
1827 | 14095.55 | 498.26 | 13597.29 | 47.32
2539 | 15362.82 | 499.17 | 14863.65 | 65.84
3003 | 16286.9 | 499.21 | 15787.69 | 4.49
3394 | 16285.96 | 499.34 | 15786.62 | 4.85
Low gain | 130 degrees | Inside
0 | 3274.91 | 494.8 | 2780.11 | 13.18
24.1 | 3156.95 | 495.05 | 2661.9 | 12.6
54.6 | 4433.7 | 494.93 | 3938.77 | 162.97
86.3 | 4428.91 | 494.86 | 3934.05 | 164.49
133 | 4584.33 | 495.22 | 4089.11 | 131.02
216 | 4472.51 | 494.66 | 3977.85 | 147.89
288 | 4639.68 | 495.03 | 4144.65 | 12.47
390 | 4624.78 | 495.26 | 4129.52 | 11.61
596 | 4529.76 | 494.65 | 4035.11 | 76.92
839 | 4551.96 | 494.72 | 4057.24 | 29.18
1007 | 4546.92 | 494.82 | 4052.1 | 54.47
1288 | 4627.92 | 494.71 | 4133.21 | 12.36
1827 | 4657.59 | 494.86 | 4162.73 | 128.06
2539 | 5010.49 | 494.85 | 4515.64 | 139.93
3003 | 5307.23 | 494.71 | 4812.52 | 13.29
3394 | 5413.25 | 495.25 | 4918 | 65.27
High gain | 130 degrees | Inside
0 | 11682.16 | 495.59 | 11186.57 | 54.19
24.1 | 11518.85 | 495.29 | 11023.56 | 52.94
54.6 | 10911.77 | 495.19 | 10416.58 | 48.49
86.3 | 10948.19 | 495.1 | 10453.09 | 52.88
133 | 11375.19 | 495.08 | 10880.11 | 52.79
216 | 16292.25 | 497.78 | 15794.47 | 4.46
288 | 16248.98 | 497.77 | 15751.21 | 25.07
390 | 10966.63 | 495.05 | 10471.58 | 48.86
596 | 11050.1 | 495.07 | 10555.03 | 50.04
839 | 11327.38 | 495.28 | 10832.1 | 48.38
1007 | 11261.8 | 495.11 | 10766.69 | 47.37
1288 | 16288.52 | 498.01 | 15790.51 | 5.79
1827 | 16292.93 | 497.96 | 15794.97 | 4.02
2539 | 16290.93 | 497.92 | 15793.01 | 4.21
3003 | 16288.36 | 497.83 | 15790.53 | 4.44
3394 | 16287.29 | 498.2 | 15789.09 | 4.01
Low gain | 90 degree | Outside+Extra light source
0 | 2760.63 | 1738.57 | 1022.06 | 5.44
24.1 | 4383.12 | 1721.84 | 2661.28 | 7.61
54.6 | 4414.35 | 1754.37 | 2659.98 | 10.52
86.3 | 3764.04 | 1231.35 | 2532.69 | 9.34
133 | 3978.67 | 1401.61 | 2577.06 | 7.11
216 | 4261.31 | 1543.4 | 2717.91 | 4.87
288 | 3623.78 | 1019.83 | 2603.95 | 139.65
390 | 4704.88 | 1715.82 | 2989.06 | 7.34
596 | 4533.38 | 1504.31 | 3029.07 | 7.29
839 | 4368.61 | 1198.51 | 3170.1 | 12.74
1007 | 3891.14 | 880.34 | 3010.8 | 31.78
1288 | 4021.97 | 773.92 | 3248.05 | 9.3
1827 | 4383.71 | 750.99 | 3632.72 | 12.95
2539 | 4654.92 | 739.5 | 3915.42 | 16.8
3003 | 4827.7 | 734.3 | 4093.4 | 10.55
3394 | 4952.15 | 670.09 | 4282.06 | 12.89
High gain | 90 degree | Outside+Extra light source
0 | 10420.59 | 2125.9 | 8294.69 | 19.99
24.1 | 15126.05 | 5426.07 | 9699.98 | 31.26
54.6 | 11466.18 | 2774.03 | 8692.15 | 25.15
86.3 | 12426.55 | 3290.61 | 9135.94 | 40.29
133 | 12391.33 | 3206.38 | 9184.95 | 37.46
216 | 13655.09 | 3887.16 | 9767.93 | 34.78
288 | 12801.82 | 2887.54 | 9914.28 | 21.5
390 | 15161.09 | 4256.72 | 10904.37 | 19.86
596 | 16280.1 | 4751.6 | 11528.5 | 44.77
839 | 16283.37 | 5224.01 | 11059.36 | 13.39
1007 | 16283.49 | 5185.26 | 11098.23 | 12.15
1288 | 16284.64 | 4585.59 | 11699.05 | 27.26
1827 | 16278.45 | 4425.65 | 11852.8 | 7.62
2539 | 16275.22 | 4594.24 | 11680.98 | 32.75
3003 | 16273.77 | 4316.22 | 11957.55 | 22.05
3394 | 16271.98 | 4194.74 | 12077.24 | 27.31
Low gain | 130 degrees | Outside+Extra light source
0 | 5842.93 | 1308.15 | 4534.78 | 13.71
24.1 | 6140.47 | 1563.42 | 4577.05 | 12.95
54.6 | 6096.9 | 1557.53 | 4539.37 | 18.51
86.3 | 6032.37 | 1464.24 | 4568.13 | 12.44
133 | 6012.11 | 1510.36 | 4501.75 | 11.61
216 | 5724.1 | 1351.23 | 4372.87 | 11.95
288 | 5721.34 | 1408.5 | 4312.84 | 12.84
390 | 5340.19 | 1092.33 | 4247.86 | 11.9
596 | 5324.97 | 1176.53 | 4148.44 | 10.49
839 | 5701.76 | 1385.65 | 4316.11 | 10.61
1007 | 5496.68 | 1234.26 | 4262.42 | 12.67
1288 | 5621.32 | 1275.15 | 4346.17 | 8.95
1827 | 5000.32 | 809.1 | 4191.22 | 13.58
2539 | 5409.81 | 752.99 | 4656.82 | 15.47
3003 | 5326.28 | 713.88 | 4612.4 | 14.59
3394 | 5484.42 | 710.76 | 4773.66 | 14.04
High gain | 130 degrees | Outside+Extra light source
0 | 16287.33 | 1024.69 | 15262.64 | 4.73
24.1 | 16287.51 | 1010.18 | 15277.33 | 4.45
54.6 | 16285.98 | 1019.87 | 15266.11 | 4.68
86.3 | 16285.94 | 1025.72 | 15260.22 | 5.13
133 | 16288.17 | 1015.47 | 15272.7 | 4.49
216 | 16287.83 | 1001.99 | 15285.84 | 4.87
288 | 16286.4 | 1010.52 | 15275.88 | 4.82
390 | 16286.34 | 1014.15 | 15272.19 | 4.74
596 | 16286.74 | 1003.23 | 15283.51 | 4.64
839 | 16280.33 | 1012.88 | 15267.45 | 8.1
1007 | 16171 | 1000.98 | 15170.02 | 40.41
1288 | 16237.48 | 998.26 | 15239.22 | 27.72
1827 | 16288.56 | 1029.64 | 15258.92 | 5.04
2539 | 16285.8 | 1020.91 | 15264.89 | 4.97
3003 | 16284.02 | 1032.61 | 15251.41 | 4.75
3394 | 16283.05 | 1016.3 | 15266.75 | 5.59
Low gain | 90 degree | Outside+Extra light source | 12 bits
NTU | LED ON | LED OFF | LEDON -LEDOFF | Standard deviation
0 | 891.38 | 323.93 | 567.45 | 2.59
24.1 | 900.07 | 326.12 | 573.95 | 2.41
54.6 | 871.85 | 299.11 | 572.74 | 1.92
86.3 | 926.01 | 329.79 | 596.22 | 2.26
133 | 935.66 | 332.13 | 603.53 | 2.57
216 | 953.11 | 333.08 | 620.03 | 2.99
288 | 980.29 | 330.62 | 649.67 | 2.26
390 | 1074.01 | 376.83 | 697.18 | 2.49
596 | 1108.21 | 375.29 | 732.92 | 2.03
839 | 1171.44 | 380.21 | 791.23 | 2.59
1007 | 1089.57 | 305.99 | 783.58 | 2.29
1288 | 1153.19 | 321.4 | 831.79 | 3.13
1827 | 1354.47 | 367.52 | 986.95 | 2.48
2539 | 1450.37 | 375.79 | 1074.58 | 2.06
3003 | 1471.07 | 379.26 | 1091.81 | 2.99
3394 | 1517.23 | 378.11 | 1139.12 | 2.31
0 | 891.27 | 341.46 | 549.81 | 2.77
0 | 884.31 | 336.29 | 548.02 | 2.58
0 | 879.99 | 333.13 | 546.86 | 2.42
0 | 879.73 | 333.34 | 546.39 | 2.27
Low gain | 90 degree | Inside
0 | 626.98 | 124.12 | 502.86 | 2.26
24.1 | 635.8 | 123.88 | 511.92 | 2.64
54.6 | 639.76 | 123.93 | 515.83 | 2.85
86.3 | 657.05 | 124.24 | 532.81 | 2.82
133 | 664.25 | 123.98 | 540.27 | 2.6
216 | 689.36 | 124.48 | 564.88 | 2.64
288 | 714.25 | 124.06 | 590.19 | 2.6
390 | 737.66 | 123.86 | 613.8 | 2.47
596 | 780.16 | 124.08 | 656.08 | 2.72
839 | 836.41 | 124.33 | 712.08 | 2.9
1007 | 843.38 | 124.41 | 718.97 | 2.51
1288 | 890.49 | 124.41 | 766.08 | 3.6
1827 | 1022.47 | 124.3 | 898.17 | 2.87
2539 | 1096.4 | 124.43 | 971.97 | 2.86
3003 | 1139.11 | 124.34 | 1014.77 | 2.81
3394 | 1188.99 | 124.25 | 1064.74 | 2.82
0 | 629.62 | 124.17 | 505.45 | 2.87
0 | 629.39 | 124.06 | 505.33 | 2.76
0 | 629.57 | 124.11 | 505.46 | 2.7
0 | 629.64 | 123.94 | 505.7 | 3.04
54.6 NTU
Orientation of sunlight LED | Measurement LEDs | On | Off | Diff | SD
None | 1 | 380.12 | 124.95 | 255.17 | 2.21
Into white label (diffused) | 1 | 381.97 | 126.69 | 255.28 | 2.18
Diagonal | 1 | 510.36 | 269.21 | 241.15 | 2.36
Transmission | 1
None | 2 | 637.36 | 125.07 | 512.29 | 2.79
Into white label (diffused) | 2 | 638.93 | 127.04 | 511.89 | 2.47
Diagonal | 2 | 702.37 | 265.5 | 436.87 | 2.45
Transmission | 2
54.6 NTU
External light souce | On | Off | Delta | SD
Direct sunlight | 871.85 | 299.11 | 572.74 | 1.92
Storage room | 637.36 | 125.07 | 512.29 | 2.79
Synthetic sunlight using LED | 702.37 | 265.5 | 436.87 | 2.45
Photodiode is in darkness | On | Off | Delta | SD
Both connected | 188.99 | 125.97 | 63.02 | 2.61 | Add a constant light
D2 connected only | 154.26 | 126.09 | 28.17 | 2.12
D3 connected only | 149.69 | 125.78 | 23.91 | 2.33
No measurement IR LED connected | 125.95 | 125.59 | 0.36 | 2.1
Power supply directly raw 100ma each
Both off | 126.03 | 126.13 | -0.1 | 2.17
Sunlight on | 149.62 | 149.91 | -0.29 | 2.07
Measurement on | 153.04 | 153.08 | -0.04 | 1.75
Both on | 176.37 | 176.62 | -0.25 | 1.76
Power supply directly 100ma each | On | Off | Delta
Sunlight on | 149.62 | 126.03 | 23.59
Measurement on | 153.04 | 126.03 | 27.01
Both on | 176.37 | 126.03 | 50.34
Power supply directly raw 50mA each
Both off | 126.14 | 126.08 | 0.06 | 2.11
Sunlight on | 138.89 | 138.88 | 0.01 | 2.02
Measurement on | 140.09 | 140.17 | -0.08 | 2.11
Both on | 152.53 | 152.74 | -0.21 | 2.21
Power supply directly 50ma each | On | Off | Delta
Sunlight on | 138.89 | 126.14 | 12.75
Measurement on | 140.09 | 126.14 | 13.95 | 26.7
Both on | 152.53 | 126.14 | 26.39 | -0.31
Power supply directly raw 20mA each
Both off | 125.89 | 126.08 | -0.19 | 2.06
Sunlight on | 130.8 | 130.92 | -0.12 | 1.86
Measurement on | 131.29 | 131.83 | -0.54 | 2.38
Both on | 136.3 | 136.54 | -0.24 | 2.2
Power supply directly 20ma each | On | Off | Delta
Sunlight on | 130.8 | 126.08 | 4.72
Measurement on | 131.29 | 126.08 | 5.21 | 9.93
Both on | 136.3 | 126.08 | 10.22 | 0.29
Power supply directly raw 200mA each
Both off | 126.05 | 126.2 | -0.15 | 2.08
Sunlight on | 170.07 | 169.95 | 0.12 | 2.14
Measurement on | 174.65 | 174.61 | 0.04 | 2.27
Both on | 217.99 | 218.05 | -0.06 | 1.95
Power supply directly 200ma each | On | Off | Delta
Sunlight on | 170.07 | 126.05 | 44.02
Measurement on | 174.65 | 126.05 | 48.6 | 92.62
Both on | 217.99 | 126.05 | 91.94 | -0.68
Current | Both
20 | 10.22
50 | 26.39
100 | 50.34
200 | 91.94
More direct both at 10mA | [5 empty] | More direct both at 10mA | Repeat
Both off | 126.3 | 126.33 | -0.03 | 2.26 |  | 126.34 | 126.17 | 0.17 | 2.01
Direct on | 495.29 | 495.1 | 0.19 | 2.89 |  | 494.5 | 494.77 | -0.27 | 2.28
Angled on | 156.9 | 157.14 | -0.24 | 3.11 |  | 154.33 | 154.58 | -0.25 | 2.24
Both on | 527.31 | 527.61 | -0.3 | 2.59 |  | 522.83 | 523.17 | -0.34 | 2.43
More direct both at 10mA | On | Off | Delta
Direct | 495.29 | 126.33 | 368.96 |  |  | 494.5 | 126.17 | 368.33
Angled | 156.9 | 126.33 | 30.57 | 399.53 |  | 154.33 | 126.17 | 28.16 | 396.49
Both on | 527.31 | 126.33 | 400.98 | 1.45 |  | 522.83 | 126.17 | 396.66 | 0.17
More direct both at 20mA | [5 empty] | More direct both at 20mA | Repeat
Both off | 126.3 | 126.33 | -0.03 | 2.26 |  | 126 | 126.26 | -0.26 | 2.14
Direct on | 651.56 | 651.94 | -0.38 | 2.54 |  | 700.37 | 700.66 | -0.29 | 2.51
Angled on | 178.21 | 178.44 | -0.23 | 2.1 |  | 167.69 | 167.91 | -0.22 | 1.87
Both on | 707.83 | 707.49 | 0.34 | 2.39 |  | 740.77 | 740.87 | -0.1 | 2.54
More direct both at 20mA | On | Off | Delta
Direct | 651.56 | 126.33 | 525.23 |  |  | 700.37 | 126.26 | 574.11
Angled | 178.21 | 126.33 | 51.88 | 577.11 |  | 167.69 | 126.26 | 41.43 | 615.54
Both on | 707.83 | 126.33 | 581.5 | 4.39 |  | 740.77 | 126.26 | 614.51 | -1.03
More direct both at 50mA | [5 empty] | More direct both at 50mA | Repeat
Both off | 125.84 | 125.54 | 0.3 | 2.02 |  | 125.89 | 126.11 | -0.22 | 2.23
Direct on | 1774.36 | 1773.37 | 0.99 | 3.74 |  | 1872.79 | 1873.38 | -0.59 | 2.83
Angled on | 244.79 | 244.63 | 0.16 | 2.16 |  | 235.22 | 235.61 | -0.39 | 2.28
Both on | 1893.22 | 1892.28 | 0.94 | 3.46 |  | 1985.23 | 1985.57 | -0.34 | 2.86
More direct both at 50mA | On | Off | Delta
Direct | 1774.36 | 125.54 | 1648.82 |  |  | 1872.79 | 126.11 | 1746.68
Angled | 244.79 | 125.54 | 119.25 | 1768.07 |  | 235.22 | 126.11 | 109.11 | 1855.79
Both on | 1893.22 | 125.54 | 1767.68 | -0.39 |  | 1985.23 | 126.11 | 1859.12 | 3.33
More direct both at 100mA | [5 empty] | More direct both at 100mA | Repeat
Both off | 125.89 | 125.86 | 0.03 | 2.19 |  | 126.2 | 125.95 | 0.25 | 2.21
Direct on | 3354.59 | 3352.8 | 1.79 | 4.74 |  | 3449.13 | 3449.78 | -0.65 | 3.85
Angled on | 333.69 | 333.56 | 0.13 | 2.16 |  | 337.43 | 337.54 | -0.11 | 2.07
Both on | 3568.16 | 3569.33 | -1.17 | 3.76 |  | 3678.58 | 3680.18 | -1.6 | 4.09
More direct both at 100mA | On | Off | Delta
Direct | 3354.59 | 125.86 | 3228.73 |  |  | 3449.13 | 125.95 | 3323.18
Angled | 333.69 | 125.86 | 207.83 | 3436.56 |  | 337.43 | 125.95 | 211.48 | 3534.66
Both on | 3568.16 | 125.86 | 3442.3 | 5.74 |  | 3678.58 | 125.95 | 3552.63 | 17.97
Current | Direct | Angled | Both
10 | 368.645 | 29.365 | 398.82
20 | 549.67 | 46.655 | 598.005
50 | 1697.75 | 114.18 | 1813.4
100 | 3275.955 | 209.655 | 3497.465
Repeat of above but use same holes | 90 degree
Current mA
0 | 578.93 | 126.16 | 452.77 | 2.18 | -5.72 | -6.7 | 0.98
10 | 586.95 | 133.88 | 453.07 | 2.14 | -4.421 | -5.237 | 0.8336
20 | 593.75 | 140.38 | 453.37 | 2.04 | -4.342 | -4.994 | 0.6872
30 | 603.81 | 150.16 | 453.65 | 2.03 | -1.003 | -1.471 | 0.5208
40 | 612.47 | 158.57 | 453.9 | 2.19 | 0.936 | 0.682 | 0.3244
50 | 615.07 | 161.87 | 453.2 | 2.12 | -3.185 | -2.275 | -0.822
60 | 625.53 | 171.84 | 453.69 | 2.1 | 0.554 | 1.438 | -0.7784
70 | 631.47 | 177.59 | 453.88 | 2.09 | -0.227 | 0.931 | -1.0348
80 | 642.56 | 187.74 | 454.82 | 2.85 | 4.142 | 4.824 | -0.5412
90 | 648.65 | 192.63 | 456.02 | 1.94 | 3.511 | 3.457 | 0.2124
100 | 652.26 | 195.6 | 456.66 | 2.05 | 0.4 | 0.17 | 0.406
110 | 657.77 | 201.36 | 456.41 | 2.17 | -0.811 | -0.327 | -0.2904
120 | 666.13 | 208.95 | 457.18 | 1.87 | 0.828 | 1.006 | 0.0332
130 | 678.45 | 220.27 | 458.18 | 2.06 | 6.427 | 6.069 | 0.5868
140 | 679.99 | 221.73 | 458.26 | 2.3 | 1.246 | 1.272 | 0.2204
150 | 687.62 | 228.76 | 458.86 | 2.07 | 2.155 | 2.045 | 0.374
160 | 693.28 | 234.53 | 458.75 | 1.96 | 1.094 | 1.558 | -0.1824
170 | 700.38 | 241.54 | 458.84 | 2.06 | 1.473 | 2.311 | -0.5388
180 | 711.57 | 251.9 | 459.67 | 1.89 | 5.942 | 6.414 | -0.1552
190 | 716.59 | 256.2 | 460.39 | 2.06 | 4.241 | 4.457 | 0.1184
200 | 720.1 | 259.61 | 460.49 | 1.89 | 1.03 | 1.61 | -0.228
210 | 725.62 | 263.69 | 461.93 | 2.13 | -0.171 | -0.567 | 0.7656
220 | 733.48 | 270.96 | 462.52 | 2.06 | 0.968 | 0.446 | 0.9092
230 | 739.14 | 275.75 | 463.39 | 2.17 | -0.093 | -1.021 | 1.3328
240 | 746.88 | 282.96 | 463.92 | 1.94 | 0.926 | -0.068 | 1.4164
250 | 755.52 | 291.56 | 463.96 | 2.11 | 2.845 | 2.275 | 1.01
260 | 755.63 | 291.21 | 464.42 | 2.05 | -3.766 | -4.332 | 1.0236
270 | 765.44 | 301.23 | 464.21 | 2.22 | -0.677 | -0.569 | 0.3672
280 | 770.91 | 306.08 | 464.83 | 2.22 | -1.928 | -1.976 | 0.5408
290 | 775.13 | 310.56 | 464.57 | 2.21 | -4.429 | -3.753 | -0.1656
300 | 778.4 | 313.09 | 465.31 | 1.98 | -7.88 | -7.48 | 0.128
Current mA | Through PD
0 | 602.49 | 126.79 | 475.7 | 2.14 | -35.48 | -33.91 | -1.57
10 | 632.31 | 155.29 | 477.02 | 2.45 | -33.206 | -32.511 | -0.695
20 | 681.64 | 202.38 | 479.26 | 1.98 | -11.422 | -12.522 | 1.1
30 | 687.84 | 208.3 | 479.54 | 2.14 | -32.768 | -33.703 | 0.935
40 | 747.01 | 266.86 | 480.15 | 2.08 | -1.144 | -2.244 | 1.1
50 | 765.44 | 285.33 | 480.11 | 2.02 | -10.26 | -10.875 | 0.615
60 | 786.94 | 306.74 | 480.2 | 2.29 | -16.306 | -16.566 | 0.26
70 | 831.44 | 350.34 | 481.1 | 2.3 | 0.648 | -0.067 | 0.715
80 | 880.27 | 399.86 | 480.41 | 2.08 | 21.932 | 22.352 | -0.42
90 | 899.44 | 418.48 | 480.96 | 2.83 | 13.556 | 13.871 | -0.315
100 | 916.76 | 434.76 | 482 | 2.29 | 3.33 | 3.05 | 0.28
110 | 950.22 | 467.54 | 482.68 | 2.22 | 9.244 | 8.729 | 0.515
120 | 978.26 | 495.54 | 482.72 | 2.13 | 9.738 | 9.628 | 0.11
130 | 1020.98 | 537.72 | 483.26 | 2.22 | 24.912 | 24.707 | 0.205
140 | 1056.57 | 573.16 | 483.41 | 2.41 | 32.956 | 33.046 | -0.09
150 | 1070.26 | 586.66 | 483.6 | 2.47 | 19.1 | 19.445 | -0.345
160 | 1091.96 | 608.43 | 483.53 | 2.25 | 13.254 | 14.114 | -0.86
170 | 1135.71 | 653.44 | 482.27 | 2.41 | 29.458 | 32.023 | -2.565
180 | 1168.66 | 684.87 | 483.79 | 2.22 | 34.862 | 36.352 | -1.49
190 | 1181.72 | 697.43 | 484.29 | 2.26 | 20.376 | 21.811 | -1.435
200 | 1214.45 | 729.56 | 484.89 | 2.18 | 25.56 | 26.84 | -1.28
210 | 1238.2 | 750.74 | 487.46 | 2.15 | 21.764 | 20.919 | 0.845
220 | 1241.32 | 753.42 | 487.9 | 2.13 | -2.662 | -3.502 | 0.84
230 | 1276.64 | 788.12 | 488.52 | 2.33 | 5.112 | 4.097 | 1.015
240 | 1302.39 | 813.37 | 489.02 | 2.22 | 3.316 | 2.246 | 1.07
250 | 1321.55 | 831.95 | 489.6 | 2.35 | -5.07 | -6.275 | 1.205
260 | 1340.17 | 850.65 | 489.52 | 2.32 | -13.996 | -14.676 | 0.68
270 | 1369.53 | 879.7 | 489.83 | 2.2 | -12.182 | -12.727 | 0.545
280 | 1379.53 | 889.24 | 490.29 | 2.18 | -29.728 | -30.288 | 0.56
290 | 1391.5 | 902.54 | 488.96 | 2.12 | -45.304 | -44.089 | -1.215
300 | 1424.9 | 934.53 | 490.37 | 2.23 | -39.45 | -39.2 | -0.25
The two relationship have some sort of constant offset. Should be possible to construct a 2D surface plot

## Correction factor
Low gain | 90 degree | Outside+Extra light source | 12 bits
NTU | LED ON | LED OFF | LEDON -LEDOFF | Standard deviation
0 | 891.38 | 323.93 | 567.45 | 2.59
24.1 | 900.07 | 326.12 | 573.95 | 2.41
54.6 | 871.85 | 299.11 | 572.74 | 1.92
86.3 | 926.01 | 329.79 | 596.22 | 2.26
133 | 935.66 | 332.13 | 603.53 | 2.57
216 | 953.11 | 333.08 | 620.03 | 2.99
288 | 980.29 | 330.62 | 649.67 | 2.26
390 | 1074.01 | 376.83 | 697.18 | 2.49
596 | 1108.21 | 375.29 | 732.92 | 2.03
839 | 1171.44 | 380.21 | 791.23 | 2.59
1007 | 1089.57 | 305.99 | 783.58 | 2.29
1288 | 1153.19 | 321.4 | 831.79 | 3.13
1827 | 1354.47 | 367.52 | 986.95 | 2.48
2539 | 1450.37 | 375.79 | 1074.58 | 2.06
3003 | 1471.07 | 379.26 | 1091.81 | 2.99
3394 | 1517.23 | 378.11 | 1139.12 | 2.31
0 | 891.27 | 341.46 | 549.81 | 2.77
0 | 884.31 | 336.29 | 548.02 | 2.58
0 | 879.99 | 333.13 | 546.86 | 2.42
0 | 879.73 | 333.34 | 546.39 | 2.27
Low gain | 90 degree | Inside
0 | 626.98 | 124.12 | 502.86 | 2.26
24.1 | 635.8 | 123.88 | 511.92 | 2.64
54.6 | 639.76 | 123.93 | 515.83 | 2.85
86.3 | 657.05 | 124.24 | 532.81 | 2.82
133 | 664.25 | 123.98 | 540.27 | 2.6
216 | 689.36 | 124.48 | 564.88 | 2.64
288 | 714.25 | 124.06 | 590.19 | 2.6
390 | 737.66 | 123.86 | 613.8 | 2.47
596 | 780.16 | 124.08 | 656.08 | 2.72
839 | 836.41 | 124.33 | 712.08 | 2.9
1007 | 843.38 | 124.41 | 718.97 | 2.51
1288 | 890.49 | 124.41 | 766.08 | 3.6
1827 | 1022.47 | 124.3 | 898.17 | 2.87
2539 | 1096.4 | 124.43 | 971.97 | 2.86
3003 | 1139.11 | 124.34 | 1014.77 | 2.81
3394 | 1188.99 | 124.25 | 1064.74 | 2.82
0 | 629.62 | 124.17 | 505.45 | 2.87
0 | 629.39 | 124.06 | 505.33 | 2.76
0 | 629.57 | 124.11 | 505.46 | 2.7
0 | 629.64 | 123.94 | 505.7 | 3.04
Repeat of above but use same holes | 90 degree
Current mA
0 | 578.93 | 126.16 | 452.77 | 2.18
10 | 586.95 | 133.88 | 453.07 | 2.14
20 | 593.75 | 140.38 | 453.37 | 2.04
30 | 603.81 | 150.16 | 453.65 | 2.03
40 | 612.47 | 158.57 | 453.9 | 2.19
50 | 615.07 | 161.87 | 453.2 | 2.12
60 | 625.53 | 171.84 | 453.69 | 2.1
70 | 631.47 | 177.59 | 453.88 | 2.09
80 | 642.56 | 187.74 | 454.82 | 2.85
90 | 648.65 | 192.63 | 456.02 | 1.94
100 | 652.26 | 195.6 | 456.66 | 2.05
110 | 657.77 | 201.36 | 456.41 | 2.17
120 | 666.13 | 208.95 | 457.18 | 1.87
130 | 678.45 | 220.27 | 458.18 | 2.06
140 | 679.99 | 221.73 | 458.26 | 2.3
150 | 687.62 | 228.76 | 458.86 | 2.07
160 | 693.28 | 234.53 | 458.75 | 1.96
170 | 700.38 | 241.54 | 458.84 | 2.06
180 | 711.57 | 251.9 | 459.67 | 1.89
190 | 716.59 | 256.2 | 460.39 | 2.06
200 | 720.1 | 259.61 | 460.49 | 1.89
210 | 725.62 | 263.69 | 461.93 | 2.13
220 | 733.48 | 270.96 | 462.52 | 2.06
230 | 739.14 | 275.75 | 463.39 | 2.17
240 | 746.88 | 282.96 | 463.92 | 1.94
250 | 755.52 | 291.56 | 463.96 | 2.11
260 | 755.63 | 291.21 | 464.42 | 2.05
270 | 765.44 | 301.23 | 464.21 | 2.22
280 | 770.91 | 306.08 | 464.83 | 2.22
290 | 775.13 | 310.56 | 464.57 | 2.21
300 | 778.4 | 313.09 | 465.31 | 1.98
Current mA | Through PD
0 | 602.49 | 126.79 | 475.7 | 2.14
10 | 632.31 | 155.29 | 477.02 | 2.45
20 | 681.64 | 202.38 | 479.26 | 1.98
30 | 687.84 | 208.3 | 479.54 | 2.14
40 | 747.01 | 266.86 | 480.15 | 2.08
50 | 765.44 | 285.33 | 480.11 | 2.02
60 | 786.94 | 306.74 | 480.2 | 2.29
70 | 831.44 | 350.34 | 481.1 | 2.3
80 | 880.27 | 399.86 | 480.41 | 2.08
90 | 899.44 | 418.48 | 480.96 | 2.83
100 | 916.76 | 434.76 | 482 | 2.29
110 | 950.22 | 467.54 | 482.68 | 2.22
120 | 978.26 | 495.54 | 482.72 | 2.13
130 | 1020.98 | 537.72 | 483.26 | 2.22
140 | 1056.57 | 573.16 | 483.41 | 2.41
150 | 1070.26 | 586.66 | 483.6 | 2.47
160 | 1091.96 | 608.43 | 483.53 | 2.25
170 | 1135.71 | 653.44 | 482.27 | 2.41
180 | 1168.66 | 684.87 | 483.79 | 2.22
190 | 1181.72 | 697.43 | 484.29 | 2.26
200 | 1214.45 | 729.56 | 484.89 | 2.18
210 | 1238.2 | 750.74 | 487.46 | 2.15
220 | 1241.32 | 753.42 | 487.9 | 2.13
230 | 1276.64 | 788.12 | 488.52 | 2.33
240 | 1302.39 | 813.37 | 489.02 | 2.22
250 | 1321.55 | 831.95 | 489.6 | 2.35
260 | 1340.17 | 850.65 | 489.52 | 2.32
270 | 1369.53 | 879.7 | 489.83 | 2.2
280 | 1379.53 | 889.24 | 490.29 | 2.18
290 | 1391.5 | 902.54 | 488.96 | 2.12
300 | 1424.9 | 934.53 | 490.37 | 2.23

## Grounding
Turbidity | Sunlight | Current | Signal
54.6 | Off | Lowered | Default | 254.66 | 127.4 | 127.26 | 2.14
54.6 | 200mA | Lowered | Default | 298.95 | 171.16 | 127.79 | 2.24
54.6 | Off | Default | Default | 309.77 | 124.32 | 185.45 | 3.53
54.6 | 200mA | Default | Default | 354.79 | 165.55 | 189.24 | 3.56

## Sunlight error
| On | Off | Diff | SD |  | On | Off | Diff | SD | On | Off | Diff | SD
Baseline | 1130.78 | 130.31 | 1000.47 | 10.09
Through photodiode | 1306.24 | 313.41 | 992.83 | 9
Through LED | 1148.52 | 137.31 | 1011.21 | 10.15
Alteration | Max current and reduced delay
| On | Off | Diff | SD
Baseline | 500.91 | 118.66 | 382.25 | 3.36
Through photodiode | 657.99 | 280.11 | 377.88 | 3.81
Through LED | 512.02 | 125.08 | 386.94 | 5
Alteration | Reduce current
| On | Off | Diff | SD
Baseline | 805.69 | 114.43 | 691.26 | 3.16
Through photodiode | 961.6 | 282.09 | 679.51 | 2.66
Through LED | 819.96 | 121.51 | 698.45 | 3.43
Alteration | Higher signal voltage
| On | Off | Diff | SD
Baseline | 2328.57 | 1502.57 | 826 | 2.63
Through photodiode | 2480.2 | 1661 | 819.2 | 2.24
Through LED | 2361.57 | 1507.41 | 854.16 | 2.36
Alteration | Increased off delay from 2 to 2000us
| On | Off | Diff | SD
Baseline | 2016.71 | 1504.99 | 511.72 | 2.24
Through photodiode | 2168.23 | 1661.6 | 506.63 | 2.71
Through LED | 2021.17 | 1505.68 | 515.49 | 2.03
With 54.6 NTU outside
| On | Off | Diff | SD
Baseline | 1989.51 | 1493.84 | 495.67 | 2.47
Through photodiode | 2178.07 | 1670.43 | 507.64 | 2.98
Through LED | 2018.89 | 1497.23 | 521.66 | 2.35
Alteration | With 54.6 NTU storage room
| On | Off | Diff | SD
Baseline | 704.91 | 119.4 | 585.51 | 2.44
Through photodiode | 858.89 | 282.86 | 576.03 | 2.67
Through LED | 733.52 | 122.77 | 610.75 | 2.87
| Removed all modifications with 54.6 NTU |  |  |  | next to window. Increased delay to 100us.
| On | Off | Diff | SD
Baseline | 774.33 | 126.52 | 647.81 | 1.98
Through photodiode | 950.56 | 323.22 | 627.34 | 2.28 | digitalWriteFast(OP_AMP_LED_PWR, LOW);
Through LED | 789.93 | 129.28 | 660.65 | 2.79 | delayMicroseconds(50);
Alteration | 54.6 NTU |  |  |  | digitalWriteFast(OP_AMP_LED_PWR, HIGH);
| On | Off | Diff | SD | delayMicroseconds(50);
Baseline | 1372.49 | 126.77 | 1245.72 | 2.65
Through photodiode | 1536.13 | 312.79 | 1223.34 | 2.65
Through LED | 1397.58 | 134.62 | 1262.96 | 2.86
Alteration | Air

## Deviation
12 Bit
Sample_len=2 | Sample_len=10 | Sample_len=100 | Delay time = 20000 | Sample_len=2 | Sample_len=10 | Sample_len=100 | Delay time = 20000
547 | 541 | 506 | 531 | 3.4085 | 3.36 | 3.7337 | 4.2829
546 | 531 | 503 | 526
550 | 538 | 503 | 533
549 | 536 | 505 | 529
546 | 541 | 505 | 526
542 | 536 | 506 | 531
547 | 531 | 501 | 525
546 | 535 | 502 | 525
543 | 537 | 496 | 528
544 | 535 | 501 | 524
545 | 536 | 500 | 530
541 | 531 | 503 | 521
548 | 536 | 499 | 523
544 | 537 | 499 | 525
542 | 532 | 500 | 530
548 | 532 | 499 | 527
547 | 537 | 504 | 520
546 | 538 | 507 | 523
548 | 533 | 497 | 524
539 | 531 | 503 | 524
547 | 534 | 500 | 527
551 | 537 | 503 | 521
555 | 532 | 504 | 524
549 | 530 | 503 | 526
552 | 539 | 501 | 528
546 | 540 | 494 | 526
552 | 537 | 503 | 528
552 | 535 | 494 | 530
544 | 536 | 506 | 526
548 | 535 | 507 | 526
541 | 533 | 505 | 520
546 | 537 | 496 | 517
539 | 538 | 495 | 531
549 | 539 | 504 | 524
545 | 532 | 500 | 522
547 | 540 | 500 | 533
544 | 537 | 496 | 530
549 | 534 | 498 | 522
545 | 535 | 498 | 527
547 | 536 | 496 | 523
544 | 534 | 500 | 530
548 | 534 | 497 | 523
545 | 535 | 494 | 529
547 | 536 | 495 | 519
548 | 535 | 500 | 532
550 | 537 | 498 | 530
546 | 540 | 499 | 520
540 | 531 | 500 | 526
544 | 531 | 500 | 524
548 | 536 | 503 | 518
541 | 532 | 502 | 522
541 | 533 | 501 | 530
549 | 535 | 497 | 527
550 | 535 | 498 | 530
551 | 533 | 503 | 527
552 | 535 | 497 | 525
556 | 533 | 500 | 529
550 | 537 | 498 | 516
550 | 536 | 491 | 527
555 | 529 | 501 | 525
548 | 532 | 493 | 530
541 | 532 | 501 | 523
550 | 535 | 501 | 530
551 | 534 | 499 | 524
550 | 535 | 505 | 522
545 | 531 | 494 | 527
550 | 537 | 500 | 521
547 | 534 | 502 | 528
545 | 534 | 496 | 525
550 | 538 | 498 | 527
549 | 533 | 495 | 519
540 | 536 | 504 | 524
549 | 531 | 498 | 526
539 | 538 | 498 | 520
549 | 534 | 501 | 521
543 | 530 | 500 | 523
550 | 536 | 499 | 528
547 | 533 | 497 | 526
551 | 531 | 498 | 527
550 | 533 | 495 | 520
551 | 533 | 497 | 524
547 | 535 | 497 | 526
547 | 534 | 498 | 527
548 | 537 | 496 | 520
545 | 534 | 497 | 519
544 | 532 | 498 | 521
543 | 534 | 496 | 530
548 | 532 | 496 | 523
549 | 530 | 497 | 529
540 | 529 | 498 | 528
546 | 538 | 494 | 528
545 | 532 | 498 | 530
548 | 530 | 501 | 524
539 | 534 | 498 | 520
543 | 537 | 499 | 523
548 | 528 | 496 | 522
548 | 529 | 497 | 519
544 | 536 | 493 | 520
541 | 537 | 501 | 526
551 | 532 | 494 | 526
545 | 529 | 492 | 518
546 | 534 | 499 | 523
545 | 532 | 492 | 526
545 | 534 | 493 | 524
547 | 524 | 497 | 525
542 | 533 | 496 | 517
543 | 538 | 499 | 524
542 | 532 | 499 | 524
545 | 531 | 497 | 524
546 | 535 | 496 | 526
546 | 528 | 500 | 527
542 | 531 | 499 | 520
541 | 531 | 494 | 529
546 | 530 | 499 | 524
548 | 543 | 499 | 527
548 | 532 | 495 | 523
540 | 527 | 499 | 526
543 | 535 | 502 | 525
545 | 538 | 495 | 518
550 | 530 | 496 | 526
550 | 530 | 495 | 525
547 | 533 | 495 | 523
545 | 538 | 497 | 523
550 | 533 | 496 | 521
546 | 529 | 493 | 521
551 | 533 | 499 | 523
546 | 533 | 497 | 522
541 | 532 | 499 | 521
550 | 530 | 497 | 523
543 | 539 | 496 | 523
538 | 530 | 495 | 524
545 | 528 | 497 | 523
544 | 528 | 500 | 524
548 | 539 | 495 | 527
546 | 524 | 495 | 524
550 | 532 | 497 | 524
546 | 530 | 499 | 520
545 | 531 | 497 | 524
545 | 540 | 496 | 521
547 | 532 | 497 | 525
545 | 536 | 494 | 522
550 | 531 | 495 | 527
544 | 537 | 493 | 531
547 | 533 | 497 | 520
543 | 535 | 495 | 517
542 | 532 | 495 | 520
546 | 532 | 497 | 520
552 | 531 | 495 | 517
541 | 532 | 499 | 522
548 | 535 | 497 | 522
541 | 534 | 498 | 524
545 | 527 | 500 | 523
545 | 534 | 496 | 520
547 | 529 | 497 | 525
545 | 528 | 499 | 523
547 | 532 | 492 | 526
548 | 539 | 497 | 522
544 | 536 | 496 | 525
540 | 541 | 496 | 521
542 | 537 | 496 | 520
549 | 535 | 494 | 521
536 | 536 | 494 | 520
545 | 532 | 494 | 524
549 | 532 | 496 | 525
546 | 531 | 498 | 523
540 | 532 | 495 | 528
543 | 529 | 498 | 517
545 | 528 | 491 | 519
547 | 535 | 498 | 519
542 | 537 | 497 | 521
548 | 531 | 493 | 516
545 | 532 | 496 | 519
542 | 539 | 496 | 519
542 | 531 | 493 | 525
547 | 538 | 498 | 522
545 | 531 | 498 | 521
540 | 534 | 498 | 521
546 | 532 | 498 | 522
546 | 537 | 492 | 527
542 | 530 | 497 | 518
543 | 531 | 494 | 520
546 | 534 | 497 | 522
543 | 525 | 497 | 520
539 | 533 | 493 | 521
546 | 535 | 491 | 521
546 | 526 | 495 | 524
545 | 539 | 495 | 523
554 | 532 | 490 | 525
549 | 540 | 493 | 522
543 | 535 | 493 | 520
543 | 540 | 499 | 521
552 | 533 | 491 | 521
542 | 535 | 495 | 517
539 | 534 | 496 | 521
542 | 539 | 494 | 525
539 | 529 | 493 | 518
542 | 538 | 490 | 522
543 | 534 | 494 | 519
546 | 540 | 496 | 517
548 | 530 | 492 | 520
546 | 534 | 497 | 525
543 | 536 | 496 | 520
543 | 538 | 500 | 526
547 | 531 | 493 | 519
546 | 531 | 495 | 522
550 | 532 | 493 | 520
542 | 532 | 498 | 522
546 | 530 | 492 | 521
542 | 526 | 496 | 524
547 | 530 | 494 | 524
548 | 534 | 500 | 527
549 | 532 | 495 | 522
544 | 532 | 498 | 526
545 | 535 | 495 | 519
542 | 526 | 498 | 522
541 | 534 | 496 | 525
543 | 533 | 493 | 522
542 | 534 | 493 | 526
545 | 528 | 495 | 519
545 | 530 | 498 | 526
547 | 535 | 495 | 521
543 | 529 | 498 | 519
546 | 531 | 497 | 514
542 | 528 | 496 | 524
544 | 529 | 494 | 523
543 | 532 | 493 | 520
540 | 532 | 495 | 519
542 | 531 | 493 | 518
543 | 535 | 492 | 521
542 | 530 | 493 | 517
541 | 533 | 498 | 523
538 | 530 | 494 | 518
546 | 527 | 497 | 519
545 | 534 | 490 | 521
542 | 535 | 495 | 516
547 | 532 | 496 | 519
548 | 529 | 497 | 519
542 | 529 | 498 | 521
546 | 531 | 499 | 522
546 | 530 | 495 | 525
546 | 534 | 495 | 517
544 | 536 | 495 | 526
546 | 531 | 498 | 530
540 | 530 | 494 | 517
546 | 534 | 493 | 524
547 | 530 | 495 | 521
544 | 538 | 495 | 529
550 | 533 | 496 | 518
545 | 532 | 498 | 523
540 | 530 | 502 | 520
547 | 533 | 495 | 520
549 | 537 | 490 | 519
544 | 531 | 497 | 519
535 | 534 | 495 | 524
549 | 532 | 497 | 522
540 | 531 | 495 | 529
547 | 532 | 492 | 527
546 | 533 | 497 | 521
547 | 536 | 496 | 520
546 | 534 | 497 | 520
545 | 533 | 495 | 523
541 | 534 | 493 | 523
543 | 530 | 496 | 519
544 | 531 | 493 | 518
540 | 525 | 496 | 522
541 | 535 | 492 | 520
543 | 536 | 495 | 520
541 | 534 | 489 | 516
547 | 537 | 499 | 520
546 | 535 | 496 | 522
544 | 536 | 494 | 524
544 | 533 | 492 | 525
545 | 534 | 489 | 522
540 | 536 | 496 | 522
541 | 531 | 493 | 526
538 | 529 | 496 | 526
541 | 534 | 494 | 522
540 | 531 | 492 | 528
541 | 528 | 487 | 522
547 | 535 | 491 | 518
549 | 530 | 495 | 519
539 | 537 | 493 | 520
544 | 530 | 491 | 526
539 | 529 | 494 | 519
547 | 534 | 495 | 520
541 | 533 | 497 | 523
549 | 529 | 499 | 526
548 | 535 | 494 | 523
542 | 528 | 494 | 518
538 | 530 | 495 | 519
545 | 526 | 497 | 525
536 | 528 | 499 | 525
534 | 533 | 492 | 519
532 | 534 | 493 | 518
543 | 534 | 497 | 521
544 | 529 | 498 | 517
545 | 534 | 494 | 526
547 | 534 | 494 | 520
545 | 529 | 492 | 514
541 | 529 | 495 | 524
548 | 531 | 490 | 526
544 | 535 | 494 | 521
544 | 531 | 498 | 519
542 | 530 | 497 | 516
547 | 532 | 488 | 518
543 | 535 | 492 | 527
543 | 534 | 493 | 520
546 | 532 | 493 | 519
547 | 530 | 496 | 525
539 | 531 | 498 | 527
543 | 529 | 498 | 524
548 | 533 | 492 | 519
544 | 533 | 497 | 519
545 | 533 | 490 | 525
544 | 529 | 494 | 517
551 | 532 | 493 | 519
542 | 535 | 487 | 523
544 | 533 | 495 | 527
544 | 535 | 490 | 527
545 | 526 | 494 | 520
544 | 535 | 495 | 527
545 | 531 | 494 | 524
543 | 529 | 492 | 523
544 | 533 | 492 | 520
541 | 533 | 498 | 518
542 | 533 | 492 | 523
546 | 530 | 493 | 521
544 | 535 | 501 | 524
548 | 534 | 490 | 525
543 | 532 | 490 | 526
543 | 530 | 490 | 521
547 | 525 | 496 | 524
549 | 534 | 495 | 524
542 | 528 | 496 | 519
543 | 532 | 497 | 515
544 | 529 | 490 | 520
546 | 534 | 493 | 515
543 | 535 | 492 | 523
545 | 531 | 497 | 523
543 | 533 | 489 | 520
544 | 532 | 489 | 521
535 | 534 | 494 | 519
542 | 536 | 495 | 522
546 | 530 | 493 | 524
548 | 532 | 491 | 524
541 | 536 | 488 | 515
542 | 534 | 500 | 517
547 | 536 | 497 | 518
549 | 529 | 494 | 520
542 | 532 | 495 | 525
542 | 530 | 498 | 525
547 | 533 | 493 | 514
540 | 531 | 492 | 524
542 | 529 | 495 | 524
543 | 528 | 495 | 522
548 | 529 | 493 | 527
541 | 530 | 494 | 520
543 | 535 | 494 | 520
544 | 532 | 494 | 524
540 | 533 | 493 | 522
547 | 537 | 499 | 517
540 | 530 | 493 | 522
540 | 534 | 495 | 520
547 | 530 | 492 | 521
548 | 530 | 495 | 517
541 | 530 | 495 | 516
552 | 534 | 495 | 520
548 | 528 | 491 | 527
541 | 534 | 490 | 521
546 | 536 | 490 | 522
549 | 527 | 497 | 521
547 | 524 | 499 | 520
546 | 536 | 491 | 524
543 | 530 | 497 | 520
543 | 525 | 496 | 519
545 | 534 | 492 | 520
543 | 535 | 494 | 524
545 | 533 | 492 | 525
546 | 529 | 496 | 522
542 | 528 | 497 | 520
543 | 533 | 489 | 519
543 | 529 | 498 | 527
546 | 529 | 489 | 524
544 | 532 | 494 | 517
544 | 533 | 493 | 519
540 | 534 | 498 | 517
542 | 529 | 493 | 518
540 | 523 | 496 | 524
543 | 533 | 497 | 524
541 | 530 | 493 | 516
541 | 532 | 493 | 523
540 | 532 | 490 | 522
544 | 535 | 492 | 523
540 | 533 | 496 | 523
542 | 529 | 492 | 524
545 | 529 | 495 | 523
543 | 528 | 494 | 521
539 | 533 | 496 | 519
545 | 531 | 492 | 514
540 | 535 | 492 | 522
540 | 531 | 491 | 519
538 | 525 | 491 | 521
544 | 531 | 490 | 517
551 | 531 | 492 | 526
549 | 537 | 494 | 521
543 | 532 | 489 | 523
541 | 526 | 497 | 516
544 | 529 | 494 | 518
543 | 530 | 488 | 525
547 | 532 | 495 | 523
545 | 527 | 493 | 521
542 | 535 | 488 | 521
543 | 533 | 500 | 526
545 | 530 | 491 | 522
539 | 532 | 488 | 520
541 | 532 | 492 | 521
547 | 529 | 490 | 520
547 | 529 | 493 | 523
542 | 534 | 492 | 528
542 | 532 | 494 | 519
543 | 523 | 490 | 521
547 | 531 | 496 | 521
541 | 535 | 493 | 518
539 | 531 | 491 | 527
542 | 531 | 489 | 518
541 | 535 | 492 | 520
544 | 534 | 496 | 522
542 | 529 | 492 | 525
544 | 528 | 497 | 517
542 | 531 | 493 | 518
541 | 529 | 492 | 520
543 | 531 | 493 | 522
545 | 529 | 494 | 518
540 | 537 | 491 | 519
539 | 536 | 494 | 514
539 | 532 | 488 | 524
544 | 534 | 495 | 515
542 | 533 | 493 | 520
543 | 537 | 489 | 520
543 | 531 | 493 | 523
545 | 534 | 493 | 517
542 | 527 | 492 | 525
547 | 532 | 492 | 517
539 | 531 | 492 | 523
546 | 533 | 495 | 521
544 | 532 | 492 | 522
542 | 538 | 496 | 526
538 | 531 | 492 | 517
541 | 530 | 493 | 520
544 | 527 | 498 | 518
545 | 530 | 495 | 525
544 | 531 | 491 | 521
545 | 530 | 488 | 520
541 | 528 | 499 | 519
543 | 531 | 490 | 526
537 | 530 | 490 | 525
544 | 530 | 494 | 521
547 | 538 | 494 | 522
541 | 532 | 489 | 522
543 | 527 | 490 | 525
543 | 529 | 487 | 522
546 | 535 | 496 | 523
541 | 535 | 494 | 520
542 | 530 | 496 | 519
542 | 531 | 496 | 525
547 | 536 | 493 | 519
546 | 523 | 493 | 524
538 | 532 | 496 | 528
543 | 532 | 492 | 521
545 | 527 | 495 | 522
542 | 534 | 495 | 519
540 | 528 | 489 | 519
545 | 530 | 489 | 524
544 | 534 | 486 | 519
543 | 530 | 489 | 516
541 | 529 | 493 | 522
543 | 528 | 492 | 515
541 | 531 | 492 | 518
542 | 536 | 491 | 522
541 | 528 | 493 | 516
541 | 534 | 494 | 524
543 | 535 | 493 | 519
545 | 529 | 491 | 523
541 | 527 | 491 | 521
541 | 529 | 498 | 516
540 | 534 | 492 | 522
543 | 533 | 492 | 522
543 | 534 | 496 | 525
543 | 532 | 490 | 518
544 | 538 | 493 | 522
549 | 528 | 495 | 524
540 | 532 | 489 | 525
545 | 535 | 489 | 519
541 | 536 | 493 | 517
540 | 531 | 488 | 516
541 | 530 | 494 | 521
548 | 531 | 495 | 524
546 | 526 | 494 | 527
547 | 529 | 494 | 519
546 | 531 | 495 | 522
539 | 534 | 492 | 521
545 | 530 | 491 | 527
539 | 529 | 495 | 520
548 | 534 | 493 | 519
544 | 531 | 487 | 524
540 | 534 | 497 | 519
545 | 526 | 494 | 518
545 | 527 | 492 | 522
541 | 531 | 493 | 514
542 | 533 | 495 | 518
538 | 531 | 494 | 523
538 | 527 | 492 | 520
546 | 533 | 493 | 518
545 | 531 | 485 | 514
542 | 532 | 491 | 517
543 | 529 | 488 | 519
549 | 531 | 493 | 518
542 | 534 | 497 | 525
543 | 531 | 496 | 525
542 | 533 | 494 | 521
545 | 525 | 494 | 525
539 | 530 | 489 | 515
546 | 530 | 491 | 523
544 | 523 | 489 | 525
542 | 535 | 493 | 521
541 | 530 | 493 | 517
547 | 529 | 492 | 523
545 | 530 | 492 | 519
541 | 529 | 489 | 526
545 | 532 | 495 | 517
538 | 527 | 494 | 519
538 | 528 | 495 | 523
544 | 530 | 493 | 525
548 | 535 | 494 | 520
542 | 527 | 493 | 519
537 | 534 | 495 | 518
541 | 533 | 493 | 525
544 | 531 | 493 | 525
547 | 530 | 487 | 519
544 | 529 | 494 | 517
539 | 534 | 490 | 514
540 | 534 | 493 | 519
544 | 532 | 496 | 518
546 | 529 | 490 | 518
538 | 529 | 487 | 519
546 | 534 | 491 | 525
541 | 528 | 494 | 519
546 | 530 | 493 | 520
539 | 535 | 494 | 515
544 | 527 | 491 | 521
545 | 530 | 494 | 522
546 | 529 | 490 | 526
546 | 530 | 497 | 517
545 | 528 | 490 | 525
543 | 534 | 497 | 528
546 | 529 | 491 | 522
547 | 535 | 494 | 511
541 | 536 | 492 | 520
536 | 530 | 494 | 518
544 | 526 | 492 | 519
542 | 533 | 495 | 521
542 | 533 | 493 | 520
543 | 529 | 493 | 522
549 | 530 | 495 | 515
543 | 531 | 492 | 526
544 | 529 | 489 | 523
543 | 529 | 496 | 517
538 | 531 | 487 | 521
539 | 532 | 490 | 514
543 | 531 | 494 | 525
545 | 524 | 490 | 519
541 | 530 | 491 | 518
543 | 533 | 491 | 523
543 | 535 | 492 | 521
546 | 523 | 489 | 522
542 | 530 | 490 | 519
544 | 532 | 489 | 520
543 | 534 | 495 | 515
544 | 535 | 497 | 522
543 | 527 | 495 | 519
543 | 532 | 490 | 520
544 | 526 | 497 | 524
540 | 528 | 494 | 524
545 | 527 | 495 | 519
536 | 527 | 493 | 524
543 | 530 | 499 | 525
540 | 529 | 496 | 517
546 | 527 | 493 | 521
546 | 531 | 492 | 519
547 | 531 | 490 | 520
543 | 530 | 492 | 520
550 | 528 | 495 | 518
549 | 531 | 490 | 521
545 | 530 | 495 | 520
544 | 528 | 494 | 523
542 | 531 | 495 | 523
548 | 531 | 492 | 521
540 | 526 | 490 | 517
539 | 532 | 495 | 524
542 | 528 | 491 | 525
546 | 528 | 490 | 515
547 | 532 | 491 | 518
544 | 533 | 490 | 523
543 | 531 | 494 | 522
542 | 527 | 482 | 526
545 | 528 | 496 | 520
544 | 529 | 493 | 521
543 | 532 | 484 | 517
544 | 524 | 495 | 523
537 | 529 | 491 | 523
549 | 527 | 492 | 522
537 | 529 | 487 | 523
546 | 529 | 491 | 522
543 | 532 | 490 | 518
539 | 526 | 496 | 525
546 | 527 | 494 | 519
543 | 529 | 490 | 525
544 | 530 | 493 | 523
546 | 528 | 491 | 521
541 | 531 | 494 | 518
541 | 525 | 489 | 520
542 | 533 | 496 | 520
541 | 527 | 491 | 521
544 | 532 | 491 | 521
549 | 530 | 491 | 519
543 | 530 | 491 | 525
542 | 528 | 488 | 516
543 | 527 | 491 | 519
544 | 531 | 490 | 519
547 | 529 | 490 | 517
546 | 528 | 488 | 524
543 | 530 | 491 | 520
540 | 535 | 489 | 518
545 | 524 | 495 | 517
548 | 523 | 492 | 525
546 | 527 | 492 | 521
544 | 531 | 493 | 522
541 | 528 | 488 | 520
541 | 532 | 495 | 518
540 | 525 | 486 | 520
543 | 532 | 493 | 517
543 | 535 | 491 | 524
544 | 532 | 493 | 513
545 | 532 | 496 | 514
546 | 527 | 493 | 516
545 | 531 | 493 | 524
544 | 529 | 491 | 520
538 | 527 | 497 | 515
541 | 525 | 495 | 520
538 | 532 | 492 | 530
543 | 532 | 495 | 523
544 | 527 | 494 | 515
545 | 535 | 494 | 518
545 | 536 | 492 | 517
546 | 529 | 494 | 527
546 | 531 | 491 | 518
544 | 533 | 494 | 512
540 | 532 | 494 | 514
543 | 531 | 489 | 523
541 | 525 | 489 | 527
545 | 528 | 490 | 523
541 | 527 | 486 | 520
540 | 528 | 490 | 524
542 | 531 | 497 | 516
542 | 530 | 490 | 514
541 | 529 | 494 | 512
545 | 529 | 493 | 520
540 | 534 | 492 | 523
541 | 529 | 489 | 514
542 | 529 | 490 | 513
549 | 535 | 491 | 519
538 | 533 | 497 | 525
543 | 526 | 492 | 522
543 | 525 | 489 | 514
549 | 526 | 492 | 510
550 | 526 | 491 | 523
544 | 530 | 493 | 525
546 | 530 | 490 | 524
538 | 528 | 490 | 523
547 | 523 | 496 | 517
548 | 536 | 493 | 525
543 | 532 | 487 | 524
542 | 528 | 492 | 522
541 | 529 | 491 | 516
544 | 532 | 492 | 522
546 | 532 | 484 | 521
544 | 532 | 494 | 516
543 | 526 | 487 | 521
549 | 532 | 492 | 525
540 | 531 | 492 | 521
544 | 532 | 494 | 520
540 | 531 | 492 | 519
544 | 529 | 494 | 515
545 | 528 | 493 | 520
545 | 528 | 490 | 521
541 | 536 | 491 | 519
537 | 531 | 489 | 522
543 | 526 | 491 | 519
539 | 528 | 491 | 523
546 | 536 | 495 | 520
539 | 533 | 491 | 518
543 | 532 | 494 | 515
544 | 531 | 489 | 524
542 | 530 | 494 | 528
539 | 532 | 491 | 516
547 | 529 | 492 | 515
541 | 527 | 491 | 521
541 | 529 | 491 | 516
541 | 531 | 493 | 515
543 | 527 | 492 | 517
545 | 528 | 492 | 517
541 | 531 | 494 | 518
547 | 527 | 488 | 528
538 | 533 | 488 | 520
543 | 530 | 489 | 522
543 | 531 | 491 | 516
539 | 536 | 494 | 511
543 | 526 | 491 | 511
544 | 531 | 494 | 509
544 | 531 | 493 | 517
544 | 529 | 494 | 514
546 | 531 | 488 | 523
542 | 532 | 494 | 511
542 | 533 | 497 | 517
540 | 528 | 488 | 522
539 | 527 | 490 | 525
543 | 532 | 490 | 526
545 | 533 | 490 | 520
546 | 526 | 494 | 523
551 | 532 | 491 | 510
549 | 531 | 491 | 513
538 | 533 | 492 | 518
539 | 528 | 490 | 517
543 | 526 | 491 | 509
540 | 535 | 491 | 512
538 | 526 | 493 | 509
543 | 529 | 490 | 519
544 | 526 | 492 | 524
545 | 530 | 490 | 522
545 | 531 | 493 | 517
540 | 528 | 491 | 516
542 | 527 | 492 | 515
545 | 528 | 491 | 519
541 | 525 | 492 | 524
540 | 527 | 492 | 522
545 | 529 | 493 | 515
540 | 530 | 491 | 521
544 | 536 | 493 | 520
543 | 527 | 493 | 521
541 | 526 | 496 | 511
540 | 530 | 491 | 521
544 | 531 | 489 | 515
548 | 531 | 493 | 514
542 | 532 | 491 | 514
541 | 531 | 491 | 514
542 | 528 | 488 | 522
544 | 531 | 493 | 519
545 | 529 | 493 | 520
545 | 526 | 491 | 514
547 | 535 | 493 | 513
539 | 528 | 493 | 523
544 | 528 | 486 | 524
540 | 531 | 486 | 519
537 | 533 | 490 | 516
539 | 527 | 492 | 527
542 | 533 | 491 | 512
541 | 529 | 489 | 517
541 | 537 | 487 | 512
544 | 528 | 491 | 513
543 | 529 | 491 | 523
539 | 533 | 494 | 522
544 | 531 | 494 | 522
548 | 532 | 487 | 510
540 | 530 | 490 | 523
541 | 533 | 495 | 516
547 | 532 | 487 | 525
544 | 533 | 497 | 527
542 | 531 | 495 | 518
542 | 530 | 495 | 526
543 | 530 | 488 | 522
542 | 530 | 491 | 521
539 | 529 | 490 | 523
541 | 531 | 497 | 514
542 | 528 | 493 | 517
546 | 529 | 492 | 517
541 | 532 | 490 | 521
543 | 529 | 488 | 511
543 | 530 | 492 | 518
538 | 526 | 490 | 522
542 | 533 | 488 | 515
538 | 527 | 492 | 511
536 | 531 | 489 | 516
543 | 530 | 491 | 520
546 | 532 | 491 | 517
542 | 529 | 490 | 519
539 | 530 | 490 | 521
540 | 534 | 491 | 517
544 | 527 | 488 | 525
541 | 522 | 491 | 526
549 | 539 | 491 | 519
539 | 527 | 489 | 523
546 | 528 | 492 | 519
541 | 527 | 490 | 517
544 | 531 | 488 | 519
543 | 531 | 492 | 517
543 | 531 | 490 | 521
545 | 533 | 491 | 524
540 | 529 | 493 | 521
543 | 530 | 493 | 513
541 | 531 | 493 | 517
540 | 533 | 490 | 522
542 | 530 | 489 | 515
542 | 525 | 497 | 518
543 | 531 | 493 | 519
545 | 530 | 491 | 516
541 | 531 | 489 | 521
549 | 531 | 493 | 521
543 | 533 | 495 | 519
546 | 530 | 495 | 519
544 | 528 | 494 | 515
549 | 533 | 492 | 514
537 | 529 | 489 | 519
546 | 532 | 496 | 515
546 | 527 | 492 | 517
539 | 529 | 490 | 521
541 | 531 | 492 | 520
537 | 528 | 491 | 513
545 | 530 | 497 | 518
544 | 531 | 493 | 521
540 | 532 | 489 | 516
547 | 531 | 497 | 519
543 | 527 | 493 | 527
542 | 525 | 494 | 519
544 | 526 | 489 | 524
541 | 530 | 492 | 525
541 | 533 | 493 | 520
542 | 527 | 490 | 525
539 | 529 | 492 | 517
546 | 530 | 492 | 518
551 | 529 | 488 | 518
543 | 529 | 491 | 522
545 | 534 | 494 | 524
547 | 532 | 489 | 520
545 | 529 | 488 | 519
537 | 527 | 495 | 517
544 | 533 | 491 | 517
549 | 531 | 492 | 521
542 | 526 | 495 | 521
536 | 530 | 491 | 521
547 | 526 | 492 | 519
550 | 530 | 488 | 524
543 | 529 | 494 | 518
538 | 528 | 487 | 519
540 | 532 | 494 | 515
547 | 531 | 489 | 520
547 | 524 | 488 | 515
542 | 527 | 490 | 518
544 | 531 | 491 | 521
546 | 526 | 488 | 520
540 | 525 | 490 | 519
539 | 526 | 490 | 521
543 | 532 | 491 | 513
548 | 529 | 488 | 518
545 | 527 | 493 | 513
550 | 528 | 495 | 523
544 | 530 | 491 | 525
546 | 536 | 492 | 517
549 | 528 | 489 | 518
542 | 531 | 492 | 526
542 | 532 | 494 | 516
553 | 528 | 494 | 517
543 | 535 | 492 | 514
536 | 533 | 492 | 516
534 | 530 | 489 | 523
545 | 527 | 494 | 519
545 | 530 | 490 | 517
539 | 535 | 489 | 527
542 | 529 | 495 | 526
541 | 532 | 491 | 518
546 | 527 | 492 | 515
542 | 532 | 491 | 512
545 | 533 | 490 | 520
545 | 524 | 490 | 511
543 | 529 | 495 | 511
531 | 529 | 491 | 519
547 | 530 | 491 | 514
548 | 531 | 493 | 521
548 | 532 | 492 | 513
545 | 525 | 489 | 512
544 | 529 | 490 | 513
538 | 531 | 489 | 510
542 | 533 | 488 | 525
539 | 529 | 493 | 511
545 | 529 | 491 | 509
540 | 532 | 495 | 518
543 | 529 | 493 | 515
540 | 533 | 486 | 516
542 | 534 | 487 | 516
537 | 528 | 496 | 516
543 | 532 | 491 | 515
538 | 533 | 494 | 511
543 | 530 | 489 | 520
542 | 527 | 486 | 517
541 | 533 | 490 | 509
546 | 529 | 487 | 523
544 | 529 | 489 | 520
543 | 536 | 491 | 517
541 | 529 | 492 | 519
550 | 534 | 491 | 516
544 | 527 | 488 | 519
542 | 530 | 493 | 521
542 | 529 | 495 | 517
545 | 526 | 491 | 523
535 | 532 | 487 | 522
543 | 532 | 491 | 514
541 | 530 | 495 | 513
546 | 536 | 492 | 518
544 | 532 | 488 | 513
543 | 529 | 492 | 520
546 | 541 | 493 | 519
543 | 529 | 491 | 519
543 | 533 | 490 | 523
544 | 529 | 490 | 523
541 | 531 | 491 | 517
543 | 529 | 493 | 511
541 | 530 | 495 | 521
544 | 529 | 487 | 511
542 | 531 | 493 | 524
540 | 527 | 492 | 517
549 | 534 | 494 | 517
545 | 524 | 492 | 526
544 | 533 | 491 | 533
552 | 534 | 495 | 513
541 | 534 | 485 | 510
538 | 530 | 494 | 515
536 | 528 | 486 | 513
538 | 533 | 490 | 521
542 | 533 | 493 | 522
539 | 533 | 494 | 512
542 | 526 | 490 | 519
537 | 532 | 491 | 516
548 | 536 | 494 | 519
539 | 528 | 489 | 523
549 | 530 | 487 | 515
540 | 537 | 486 | 527
543 | 535 | 487 | 530
543 | 526 | 491 | 516
541 | 528 | 490 | 520
545 | 532 | 492 | 520
539 | 531 | 485 | 523
540 | 526 | 492 | 518
540 | 530 | 487 | 524
539 | 533 | 489 | 522
546 | 528 | 490 | 514
545 | 530 | 487 | 528
543 | 532 | 493 | 517
539 | 534 | 491 | 518
545 | 529 | 490 | 514
537 | 530 | 484 | 513
546 | 530 | 491 | 516
544 | 526 | 488 | 525
545 | 531 | 494 | 511
544 | 531 | 492 | 523
542 | 534 | 496 | 520
545 | 532 | 490 | 522
541 | 527 | 496 | 521
545 | 529 | 488 | 526
542 | 532 | 490 | 519
547 | 531 | 488 | 525
544 | 528 | 492 | 525
546 | 533 | 492 | 510
544 | 528 | 493 | 515
539 | 529 | 488 | 517
543 | 532 | 492 | 518
541 | 527 | 494 | 523
540 | 531 | 491 | 522
541 | 526 | 490 | 514
546 | 534 | 489 | 520
542 | 530 | 496 | 519
547 | 530 | 492 | 512
543 | 527 | 490 | 517
544 | 525 | 489 | 520
542 | 531 | 487 | 513
542 | 529 | 492 | 515
543 | 533 | 495 | 513
545 | 528 | 489 | 522
540 | 529 | 492 | 516
545 | 529 | 496 | 523
541 | 542 | 488 | 523
546 | 528 | 488 | 517
540 | 534 | 494 | 516
537 | 528 | 492 | 521
544 | 532 | 487 | 516
545 | 535 | 489 | 519
537 | 528 | 500 | 516
546 | 532 | 492 | 524
541 | 533 | 491 | 527
542 | 530 | 486 | 513
542 | 525 | 490 | 510
544 | 531 | 489 | 516
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.
This chart isn't available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.

## Deviation 2
12 bits |  |  |  | 1000 trials | 14 bits |  |  |  | 16 bits
High delay | Low delay | Sample len | SD |  | High delay | Low delay | Sample len | SD | High delay | Low delay | Sample len | SD
2 | 20000 | 2 | 3.74 |  | 0 | 20000 | 2 | 38.36 | 0 | 20000 | 0 | 419.26
2 | 2000 | 2 | 6.02 |  | 0 | 20000 | 20 | 62.02
2 | 2000 | 2 | 6.02 |  | 0 | 20000 | 0 | 26.51
2 | 200000 | 2 | 3.36 |  | 0 | 20000 | 0 | 9.93
20 | 20000 | 2 | 3.99 |  | 0 | 20000 | 0 | 12.29
0 | 20000 | 2 | 3.53
0 | 20000 | 20 | 3.91
0 | 20000 | 0 | 152.69
0 | 20000 | 2 | 2.89

## Calcas
Burst current mA | Burst time ms | Burst time (s) | Milliamp seconds
1200 | 0.02 | 2e-05 | 0.24
| Sleep time (S) | Sleep current (ma) | Milliamp seconds
| 10 | 0.02 | 0.2
|  |  | mAs over 10 seconds | mAh | Year
|  |  | 0.44 | 0.00012222 | 385.4396
| Burst time ms | Burst time (s) | Milliamp seconds
| 100 | 0.1 | 2.2e-05
| Sleep time (S) | Sleep current (ma) | Milliamp seconds
| 9.9 | 0.011 | 0.1089
|  |  | mAs over 10 seconds | mAh | Year
|  |  | 0.10892 | 3.0256e-05 | 95.4157